Date:		6/1/15

To:		Subcommittee D02.BO

Tech Contact:	E. A. Hap Thompson

Work Item #:	WK50202

Ballot Action:	Revision of D6891

Rationale:	A new passenger car engine oil classification has been approved and this standard requires a 			complete revision.

Standard Test Method for

Evaluation of Automotive Engine Oils in the Sequence IVB Spark-Ignition Engine1

This standard is issued under the fixed designation D6891; the number immediately following the designation indicates the year of original adoption or, in the case of revision, the year of last revision. A number in parentheses indicates the year of last reapproval. A superscript epsilon (ε) indicates an editorial change since the last revision or reapproval.

INTRODUCTION

Portions of this test method are written for use by laboratories that make use of ASTM Test Monitoring Center (TMC)2 services (see Annex A1).

The TMC provides reference oils and engineering and statistical services to laboratories that desire to produce test results that are statistically similar to those produced by laboratories previously calibrated by the TMC.

In general, the Test Purchaser decides if a calibrated test stand is to be used. Organizations such as the American Chemistry Council require that a laboratory utilize the TMC services as part of their test registration process. In addition, the American Petroleum Institute and the Gear Lubricant Review Committee of the Lubricant Review Institute (SAE International) require that a laboratory use the TMC services in seeking qualification of oils against their specifications.

The advantage of using the TMC services to calibrate test stands is that the test laboratory (and hence the Test Purchaser) has an assurance that the test stand was operating at the proper level of test severity. It should also be borne in mind that results obtained in a non-calibrated test stand may not be the same as those obtained in a test stand participating in the ASTM TMC services process.

Laboratories that choose not to use the TMC services may simply disregard these portions.

1.  Scope

1.1 This test method measures the ability of crankcase oil to control valve-train wear for spark-ignition engines at low operating temperature conditions. This test method is designed to simulate extended engine cyclic vehicle operation. The Sequence IVB Test Method uses a Toyota 2NR-FE water cooled, 4 cycle, in-line cylinder, 1.5 liter engine. The primary result is bucket lifter wear. Secondary results include cam lobe nose wear and measurement of iron wear metal concentration in the used engine oil. Other determinations such as fuel dilution of the crankcase oil, non-ferrous wear metal concentrations, total fuel consumption, and total oil consumption, can be useful in the assessment of the validity of the test results.2

1.2  The values stated in SI units are to be regarded as standard. No other units of measurement are included in this standard.

1.2.1  Exceptions—Where there is no direct SI equivalent such as pipe fittings, tubing, NPT screw threads/diameters, or single source equipment specified.

1.3  This standard does not purport to address all of the safety concerns, if any, associated with its use. It is the responsibility of the user of this standard to establish appropriate safety and health practices and determine the applicability of regulatory limitations prior to use. See Annex AX for specific safety precautions.

2.  Referenced Documents

2.1  ASTM Standards:3

D235  Specification for Mineral Spirits (Petroleum Spirits) (Hydrocarbon Dry Cleaning Solvent)

D287  Test Method for API Gravity of Crude Petroleum and Petroleum Products (Hydrometer Method)

D323  Test Method for Vapor Pressure of Petroleum Products (Reid Method)

D381  Test Method for Gum Content in Fuels by Jet Evaporation

D445  Test Method for Kinematic Viscosity of Transparent and Opaque Liquids (and Calculation of Dynamic Viscosity)

D525  Test Method for Oxidation Stability of Gasoline (Induction Period Method)

D3525  Test Method for Gasoline Diluent in Used Gasoline Engine Oils by Gas Chromatography

D4485  Specification for Performance of Active API Service Category Engine Oils

D5185  Test Method for Multielement Determination of Used and Unused Lubricating Oils and Base Oils by Inductively Coupled Plasma Atomic Emission Spectrometry (ICP-AES)

D5844  Test Method for Evaluation of Automotive Engine Oils for Inhibition of Rusting (Sequence IID) (Withdrawn 2003)4

E230  Specification and Temperature-Electromotive Force (EMF) Tables for Standardized Thermocouples

2.2  API Standard:

API 1509  Engine Oil Licensing and Certification System5

2.3  SAE Standards:

SAE J183  Engine Oil Performance and Engine Service Classification6

SAE J254  Instrumentation and Techniques for Exhaust Gas Emissions Measurement6

2.4  ASME Standard:

B46.1  Standard for Surface Texture (Surface Roughness, Waviness, and Lay)7

2.5  JASO Standard:

M 328-95  Valve-train Wear Test Procedure for Evaluating Automobile Gasoline Engine Oils8

2.6  CEC Standard:

CEC-L-38-A-94   Peugeot TU-3M/KDX Valve-train Scuffing Wear Test9

3.  Terminology

3.1  Definitions:

3.1.1  blowby, n—that portion of the combustion products and unburned air/fuel mixture that leaks past piston rings into the engine crankcase during operation.

3.1.2  calibration test stand, n—a test stand on which the testing of reference material(s), conducted as specified in the standard, provided acceptable results.Sub. B Glossary10

3.1.2.1  Discussion—In several automotive lubricant standard test methods, the ASTM Test Monitoring Center provides testing guidance and determines acceptability.

3.1.3  reference oil, n—an oil of known performance characteristics, used as a basis for comparison.

3.1.3.1  Discussion—Reference oils are used to calibrate testing facilities, to compare the performance of other oils, or to evaluate other materials (such as seals) that interact with oils.

D5844

3.2  Definitions of Terms Specific to This Standard:

3.2.1  assessment length, n—the length of surface over which measurements are made.

3.2.2  break-in, n—initial engine operation to reach stabilization of the engine performance after new parts are installed in the engine.

3.2.3  cam lobe wear, n—the sum of the wear determined at the following locations (nose is zero location): (1) 14 cam degrees before the nose, (2) 10° before the nose, (3) 4° before the nose, (4) at the nose, (5) 4° after the nose, (6) 10° after the nose, (7) 14° after the nose.

3.2.4  cam nose wear, n—the maximum linear deviation of a worn nose profile from the unworn profile; the nose is the high lift point on the particular cam lobe.

3.2.5  flushing, n—the installation of a fresh charge of lubricant and oil filter for the purpose of running the engine to reduce and eliminate remnants of the previous oil charge.

3.2.5.1  Discussion—Flushing may be carried out in an iterated process to ensure a more thorough process of reducing previous oil remnants.

3.2.6  reference line, n—a deduced, leveled, straight line drawn on the profilometer graph, from the front unworn average edge of a cam lobe to the rear unworn average edge of that cam lobe.

3.2.7  valve-train, n—a mechanical engine subsystem comprised of the camshaft, the rocker arms, hydraulic lash adjusters, the poppet valves, and valve-springs.

3.2.8  wavenesstotal, n—the maximum excursion of the worn surface as graphically measured normal to the reference line.

4.  Summary of Test Method

4.1  Test Numbering Scheme—Use the test numbering scheme shown below:

AAAAA–BBBBB–CCCCC

AAAAA represents the stand number. BBBBB represents the number of tests since the last calibration test on that stand. CCCCC represents the total number of Sequence IVB tests conducted on that stand. For example, 6-10-175 represents the 175th Sequence IVB test conducted on test stand 6 and the tenth test since the last calibration test. Consecutively number all tests. Number the stand calibration tests beginning with zero for the BBBBB field. Multiple-length Sequence IVB tests are multiple runs for test numbering purposes, such as double-length tests which are counted as two runs and triple-length tests which are counted as three runs. For example, if test 1-3-28 is a doubled-length test, number the next test conducted on that stand 1-5-30.

4.2  Test Engine—This procedure uses a Toyota 2NR-FE water cooled, 4 cycle, in-line four cylinder, 1.5 liter engine as the test apparatus. The engine incorporates dual overhead camshafts, four valves per cylinder (2 intake; 2 exhaust), and a direct acting mechanical bucket lifter valve-train design. The critical test parts (camshafts, direct acting mechanical bucket lifters) are replaced prior to each test. A 95 min run-in schedule, followed by a 100 h aging schedule, for Silicon (Si) pacification, is conducted whenever the long block or cylinder head are replaced with new components.

4.3  Test Stand—The complete test stand is available from the supplier listed in Annex AX (TBD).

4.4  Test Sequence—After an engine run-in and aging schedule, or after the completion of a previous test, install new test camshafts and bucket lifters. Perform four engine flushes, using fresh oil charges for each flush. After completing the four flushes, drain the used oil from the 4th flush, and weigh and install the fresh test oil charge. Run the test for a total of 200 h, with no scheduled shutdowns. A single test cycle is composed of two 7-s steady-state stages separated by 8-s linear transitions. This test cycle (two steady-state stages and two linear transitions) is repeated 24,000 times.

4.5  Analyses Conducted—At the completion of the test, the camshaft lobes are measured for heel-to-toe wear and the bucket lifters are measured for maximum wear (z diff), area loss, volume loss and mass loss.  Use these measurements to determine the summation, average, minimum and maximum wear for the intake and exhaust bucket lifters and the intake and exhaust camshaft lobes. Determine the oil consumption by calculating the difference between the mass of the used drain oil and the mass of the engine’s initial oil charge. Analyze the end of test used oil for wear metals, fuel dilution, kinematic viscosity at 40°C, total acid number, total base number and oxidation and nitration by FTIR. Retain a final drain sample of 1 L for a minimum of 90 days. Retain the camshafts and bucket lifters for a minimum of 6 months.

5.  Significance and Use

5.1  This test method was developed to evaluate automotive lubricant’s effect on controlling camshaft lobe and bucket lifter wear for overhead camshaft engines with direct acting bucket lifters.

Note 1—This test method may be used for engine oil specifications, such as Specification D4485, API 1509, SAE J183, and ILSC GF 3.

Note 2—Coordination with the ASTM Committee D02, Subcommittee B, Sequence IVB Surveillance Panel is a prerequisite to the use of any equivalent apparatus. Figures are provided throughout the test method to suggest appropriate design details and depict some of the required apparatus.

6.  Apparatus

Note 3—Coordination with the ASTM Committee D02, Subcommittee B, Sequence IVA Surveillance Panel is a prerequisite to the use of any equivalent apparatus. However, the intent is to permit reasonable adaptation of existing laboratory facilities and equipment. Figures are provided throughout the test method to suggest appropriate design details and depict some of the required apparatus.

6.1  Test Engine—This test method uses a fired 1994 model Nissan KA24E, in-line 4-cylinder, 4-cycle, water-cooled, port fuel-injected gasoline engine with a displacement of 2.389 L.11 ,12 See Annex AX for a parts lists. Nominal oil sump volume is 3.5 L. The cylinder block is constructed of cast iron, while the cylinder head is aluminum. The engine features a single overhead camshaft with sliding follower rocker arms, with two intake valves and one exhaust valve per cylinder, and hydraulic lash adjusters. The camshaft is not phosphate-coated or lubrited. The rocker arm contact pad material is powdered metal. The engine compression ratio is 8.6 to 1. Rate the engine at 198 N·m torque at 4400 r/min. The ignition timing and multi-port fuel injection system is ECM. Fuel the engine with a specially blended, non-detergent unleaded reference gasoline. Make the EGR non-operable.

6.1.1  Engine Buildup and Measurement Area—The ambient atmosphere of the engine buildup and measurement areas shall be reasonably free of contaminants and maintained at a uniform temperature. Maintain the specific humidity at a uniform level to prevent the accumulation of rust on engine parts. Use uniform temperatures to ensure repeatable dimensional measurements. Use a sensitive surface profilometer instrument to measure the wear of the cam lobes, and place the profilometer on a base-plate free of external vibrations.

6.1.2  Engine Operating Area—The laboratory ambient atmosphere shall be reasonably free of contaminants and general wind currents, especially if and when the valve-train parts are installed while the engine remains in the operating area. The temperature and humidity level of the operating area is not specified.

6.1.3  Parts Cleaning Area—This test method does not specify the ambient atmosphere of the parts cleaning area (Warning—Use adequate ventilation in areas while using solvents and cleansers).

6.2  External Engine Modifications—Modify the test engine for the valve-train wear test. Make the exhaust gas recirculation non-operable. Disable the swirl control actuator. Disable the fast idle system and the auxiliary air control (AAC) valve. Replace the engine coolant temperature sensor by a fixed resistor. Modify the engine water-pump to incorporate an external electric-driven water-pump. Do not use the water-pump fan blade and cooling radiator. Remove the alternator. Install an oil cooler (water-to-oil heat exchanger) at the oil filter housing, as shown in Annex AX. Modify the engine wiring harness. Install fittings for various temperature and pressure measurements as required by the test method. Place the Nissan production rocker cover with a specially manufactured aluminum jacketed rocker cover. Route the engine coolant through this jacket. Install a fitting in the front engine cover to allow a portion of the crankcase ventilation air to bypass the rocker cover.

6.2.1  Non-Operable EGR—This test method does not use an EGR valve. Cover the EGR port with the supplied 3 mm thickness block-off (blind) plate (see Annex AX). Remove the hose from the exhaust manifold to the EGR. Plug the EGR supply port in the rear of the exhaust manifold with a pipe fitting.

6.2.2  Swirl Control Actuator—Disable the swirl control actuator by removing the harness connector and vacuum line. Plug the vacuum line source.

6.2.3  Fast Idle Disabling—To disable the fast idle system, remove the fast idle cam on the throttle body.

6.2.4  Engine Coolant Temperature Sensor—Substitute the variable input of the coolant temperature sensor to the ECM at the wiring harness of the ECM with a fixed resistance of 300 Ω.

6.2.5  Utility Engine Water-pump—Modify the engine water-pump shown in Fig. 1 to serve as a dummy housing on the engine, and use an electric motor-driven, external water pump for this test.

FIG. 1  Modified Water Pump

6.2.5.1  Support two surfaces, 180° apart, of the underside (non-machined surface) of the 77 mm diameter steel hub. Leave the shaft, body, and impeller free to be pressed out of the supported hub.

6.2.5.2  Using a press punch rod with the approximate diameter of 14 mm, press the shaft out of the hub.

6.2.5.3  Locate the copper wire clip in the slot on the side of the aluminum alloy pump body. Remove the U-shaped wire clip by pulling perpendicular to the longitudinal axis of the water-pump shaft.

6.2.5.4  Support the flat, machined face of the aluminum alloy pump body on two sides, 180° apart, leaving the impeller, bearings, seal, and shaft free to be pressed out of the aluminum alloy pump body.

6.2.5.5  Again using press punch rod with the approximate diameter of 14 mm, press the shaft, impeller, double bearing, and seal assembly out of the aluminum alloy pump body. Press in the direction of the internal cavity.

6.2.5.6  Clean and prepare the aluminum alloy pump body for contamination-free welding.

6.2.5.7  Fabricate a water pump bore plug (see Annex AX) starting at the neck of the aluminum alloy pump body towards the internal cavity. In some instances, due to manufacturing tolerances, the pump body may need to be heated to approximately 200 °C and the fabricated bore plug cooled to approximately 0 °C. This will allow easy installation of the bore plug.

6.2.5.8  Preheat the aluminum alloy pump body (with plug installed) to approximately 200 °C.

6.2.5.9  Using an argon/tungsten-inert gas welder with pedal/rheostat-operated 220 A, 4043 aluminum 3 mm filler rod, and the approximate settings of ac and high frequency, weld the base perimeter of the plug to the internal cavity of the aluminum pump body.

6.2.5.10  Allow to cool, then perform final cleaning before installation on the engine.

6.2.6  Coolant Bypass Hose—Disconnect the coolant bypass hose at the intake manifold. The connection ends are plugged to prevent bypass flow. Remove the thermostat.

6.2.7  Oil Cooler—Insert a water-to-oil heat exchanger (see Annex AX) between the engine oil filter adapter block and the oil filter, using a gasket as shown in Annex AX. See Annex AX for installation details. Plumb the water outlet to the cooler fitting and orient to the same axis as the oil filter. Orient the cooler for both water fittings to face the rear of the engine. To connect process water to the oil cooler, use flexible hoses (16 mm diameter) of approximately 500 mm length to connect process water to the oil cooler. Control the oil temperature by metering the flow of the process water outlet. A control system valve with Flow Coefficient (Cv) of 0.32 produces satisfactory control. Replace the oil cooler when it no longer remains serviceable.

6.2.8  Ignition Power Supply—Use a 15 A dc power supply to provide (13.4 to 14.2) V dc to the ECM that powers the engine ignition system (a Lambda Electronics Corporation Model No. LFS-43-15 has been found useful).13 ,12 Provide a separate power source for the starter motor circuit. Use an automotive battery equipped with a low-amperage battery charger.

6.2.9  ECM Wiring Harness Modifications—Remove the connectors and wires from the electronic control module wiring harness except those shown in Table 1.

TABLE 1 ECM Wiring Harness ModificationsA

| Connector Description | Connector Number(s) |
| --- | --- |
| Camshaft Position Sensor | 30M |
| Power Transistor | 44M |
| Distributor | 46M |
| Ignition Coil | 47M, 97M |
| Oxygen Sensor | 59M |
| Mass Air Flow Sensor | 63M |
| Engine Coolant Temperature Sensor | 65M (Install 300Ω resistor) |
| Throttle Position Sensor | 66M |
| Injectors 1–4 | 72M, 73M, 74M, 75M |
| Intake Air Temperature Sensor | 18M |
| Body Ground | 275M |
| Engine Ground | 60M, 61M |
| Connector Description | Connector Number(s) |
| Fuel Pump RelayB | 5M |
| ECCS RelayC | 6M |
| Resistor and Condenser | 40M |
| Check Connector | 208M |
| Joint Connector A | 259M |
| ECM (ECCS Control Module) | 262M |
| Fuel Pump | 2C |
| Joint Connector C | 212M (jumper hardwired) |
| Connector | 260M (jumper hardwired) |
| EGR Temperature sensor | 17M (retain, do not connect) |
| EGRC solenoid valve | 88M (retain, do not connect) |
| IACV-AAC Valve and | 64M (retain, do not connect) |
| IACV-FICD Solenoid Valve |  |
| Ground Connector | (retain, do not connect) |
| Check Engine Light | add and utilize |
| 30 A fuse holder | add and utilize |
| GroundD | add and utilize |
| Keep-Alive wire | add and utilize |
| Ignition wire | add and utilize |
| Ground wireD | add and utilize |

A See modified wiring diagram in Annex AX.

B Modify the fuel pump relay connector (5M) to provide a nominal 13 V to the fuel pump only when turning on the ignition power switch. See Annex AX for the wiring details.

C The ECCS relay uses the 6M connector. Connect it to the battery through a fusible link.

D Attach the wiring harness grounds to the front engine-lifting bracket.

6.3  Test Stand and Laboratory Equipment—This engine-dynamometer test is designed for operation using computer control instrumentation and computer data acquisition. Provide an intake air system for the precise control of engine intake air humidity, temperature, and cleanliness.

6.3.1  Computer Data Acquisition System—The procedure shown in 6.3.1.1 – 6.3.1.3 details the test stand log operational data with a computer data acquisition system using sensor configurations, and is in compliance with Data Acquisition and Control Automation II.10 Consider a test that has greater than 2 h without data acquisition on any controlled parameter to be operationally invalid.

6.3.1.1  Frequency of Logged Steady-State Data—Log the Stage I steady-state (last 45 min of stage) operational conditions every 2 min or more frequently. Log the Stage II steady-state (last 5 min of stage) operational conditions every 30 s or more frequently.

6.3.1.2  Frequency of Logged Transient Data—Define the transient time as the first 5 min following operational stage changes. Computer log and plot the cycle 5 transient data. Log the critical parameters (engine speed, torque, oil gallery temperature, coolant out temperature) once per second or higher frequency. If cycle 5 transients are beyond the procedural limits defined in 11.2.6, document and confirm the corrective action with the next available transition plot.

6.3.1.3  System Time Response for Logged Data—Do not exceed the controlled operational parameters for system time response for measurement shown in Table 2. The system time response includes the total system of sensor, transducer, analog signal attenuation, and computer digital filtering. Use single-pole type filters for attenuation.

TABLE 2 System Time Response

| Parameters | Time Response, max (one time constant) |
| --- | --- |
| Temperatures | 2.5 s |
| Pressures | 1.6 s |
| Coolant Flow | 2.5 s |
| Torque | 2.0 s |
| Speed | 1.8 s |

6.3.1.4  Quality Index—The Quality Index (QI) is an overall statistical measure of the variation from test targets of the steady-state operational controlled parameters. The Sequence IVA Surveillance Panel has chosen the QI upper and lower control limits, shown in Table 3.

| where: |
| --- |
| Xi | = | values of the parameter measured, |
| U | = | allowable upper limit of X, |
| L | = | allowable lower limit of X, and |
| n | = | number of data points used to calculate QI. |

TABLE 3 Upper and Lower Control Limits

| Parameter | L | U |
| --- | --- | --- |
| Coolant Flow | 29.8 | 30.2 |
| Coolant Out Temperature, Stage I and II | 49.81 54.81 | 50.19 55.19 |
| Exhaust Back-pressure | 103.34 | 103.66 |
| Intake Air Humidity | 10.8 | 12.2 |
| Intake Air Pressure | 0.047 | 0.053 |
| Intake Air Temperature | 31.71 | 32.29 |
| Oil Cylinder Head Temperature, Stage I and II | 48.7 58.7 | 49.3 59.3 |
| Speed, Stage I and II | 793.5 1493.5 | 806.5 1506.5 |
| Torque | 24.5 | 25.5 |
| Rocker Cover Air Flow | 9.5 | 10.5 |

Where missing data or Bad Quality Data (BQD), or both, are encountered, calculate the adjusted Quality Index (QIADJ) using the following equation:

| where: |
| --- |
| Q | = | QI calculated without missing/BQD, |
| I | = | points, |
| n | = | number of data points used to calculate QI, and |
| N | = | number of data points for a complete data set. |

If the QI calculation of a controlled parameter is less than zero, investigate the reason, assess its impact on test operational validity, and document such finding in the final test report. For calibration tests, review the operational validity assessment with the TMC. Annex A2 describes calibration procedures using the TMC reference oil, including their storage and conditions of use, the conducting of tests and the reporting of results.

6.3.2  Test Stand Configuration—Mount the engine on the test stand similar to its vehicle orientation (tilted up 5.5° in front; sideways 10° up on intake manifold side; bottom of oil sump horizontal). This orientation is important to the return flow of oil in the cylinder head, and ensures reproducible oil levels. Directly couple the engine flywheel to an eddy-current dynamometer through a driveshaft. The driveshaft design shall minimize vibration at the test operating conditions. The dynamometer system shall have inertia of (0.75 ± 0.15) kg·m2 to ensure satisfactory control of engine speed at 800 r/min, stable air-to-fuel ratio control, and enable reproducible transient control of engine speed and torque during stage changes. Do not use hydraulic type dynamometers, as they exhibit residual torques at low speed operation. Do not use the engine to drive any external engine accessory. Recommend the area above and to the left of the rocker arm cover be left unobstructed to allow for easier on-site replacement of valve-train wear parts while the engine rests on the test stand. See Annex AX for Safety Precautions.

6.3.3  Dynamometer Speed and Torque Control System—To improve laboratory reproducibility for transient control of engine speed and torque, the driveline system inertia, excluding engine, shall be (0.75 ± 0.15) kg·m2. Control the engine power for evaluating the lubricant in a repeatable manner by:

6.3.3.1  Measuring and controlling engine speed and dynamometer torque,

6.3.3.2  Controlling exhaust absolute pressure by exhaust pipe throttling, and

6.3.3.3  Controlling the supply of intake air temperature, humidity, and pressure differential above barometer pressure.

Note 4—The dynamometer speed and torque control systems shall be capable of maintaining the steady state operating set points within the performance envelope (that is, quality index established by the industry matrix testing program).

Note 5—Two types of full closed-loop speed and torque control systems have been successfully utilized. One typical closed-loop system maintains speed by varying dynamometer excitation and maintains torque by varying the engine throttle. This arrangement may provide satisfactory steady-state control. Another closed-loop speed and torque control system maintains torque by varying dynamometer excitation and controls speed using the engine throttle. This arrangement may provide satisfactory transient control during stage changes.

6.3.4  Intake-air Supply System—The supply system shall be capable of delivering a minimum of 600 L/min (2000 L/min preferred) of conditioned and filtered air to the test engine during the 100 h test, while maintaining the intake-air parameters detailed in Annex AX. A humidifying chamber controls the specific humidity and provides a positive air pressure to an intake air supply duct. Annex AX shows a general schematic of the intake air system.

6.3.4.1  Induction Air Humidity—Measure the intake air specific humidity in the main system duct or at the test stand. If using a main system duct dew point temperature reading to calculate the specific humidity, verify the dew point periodically at the test stand. Maintain the duct surface temperature above the dew point temperature at all points downstream of the humidity measurement point to prevent condensation and loss of humidity level.

6.3.4.2  Intake Air Filtering—Use the production intake air cleaner assembly (Annex AX), with filter, at the engine. Use a snorkel adapter, functionally equivalent to that shown in Annex AX, to connect the controlled air duct to the air cleaner. Modify the top of the air cleaner assembly for the installation of the intake temperature sensor and for the intake pressure sensor line. Refer to 6.3.4.5.

6.3.4.3  Intake Air Flow—Do not measure for intake airflow.

6.3.4.4  Intake Air Temperature—For final control of the inlet air temperature, install an electric air heater strip within the air supply duct. The duct material and heater elements design shall not generate corrosion debris that could be ingested by the engine. To provide sufficient duct flow for adequate air temperature control, it is recommended that excess air be dumped just prior to the air cleaner snorkel. An air dump area of approximately 60 mm2 will provide sufficient flow without stagnation. If additional airflow is required to stabilize air temperature, it is permissible to install a nominal 10 mm bleed hole in the air filter housing. Install the inlet temperature sensor in the air cleaner, centered at the inlet to the air cleaner (see Annex AX). Attach a support brace to the air cleaner assembly mounting stud and wing nut, if vibration of the temperature sensor is a problem.

6.3.4.5  Intake Air Supply Pressure—Install a disc type valve in the controlled air system supply duct to control the engine inlet air gage pressure. Locate the sensing tube for inlet air pressure in the topside of the air cleaner assembly ((50 ± 10) mm left and (80 ± 10) mm in front of the right rear corner of the assembly). This location senses the pressure before the air enters the air cleaner element.

6.3.5  Fuel Supply System—This test method requires approximately 200 L of unleaded Haltermann KA24E Green test fuel14 ,12 per test (100 cycles). Ensure a sufficient fuel supply at the start of test to conduct the test without a shutdown. Use the production port fuel injection system, including fuel pump (see Annex AX), fuel injector rail, and fuel pressure regulator. Use recirculated fuel within the system using a non-production heat exchanger to maintain fuel temperature ranging from (15 to 30) °C. Measure fuel consumption using a mass flow meter (MicroMotion15 ,12 model D-6 is suitable). Install a fuel filter assembly (see Annex AX) upsteam of the fuel pump. Ensure proper fuel filtration to maintain precise air-fuel ratio control during the test.

6.3.5.1  Fuel Temperature—Measure fuel temperature through one of the ports in a cross fitting located in the line between the fuel pump and the fuel rail. Maintain the fuel temperature to the fuel rail below 50 °C.

6.3.5.2  Fuel Pressure—Measure the fuel pressure through one of the ports in a cross fitting located in the line between the fuel pump and the fuel rail inlet.

6.3.5.3  Fuel Flow—Install a mass fuel flow meter for measuring the fuel consumption rate in the fuel supply system, prior to the fuel recirculating loop. A MicroMotion model D-6 fuel flow meter has been found to be suitable.

6.3.6  Exhaust System—Use a production cast iron exhaust manifold, without insulation, for the test.

6.3.6.1  Plug the rear of the manifold (EGR supply) with a pipe fitting. Do not use an EGR for this test.

6.3.6.2  Use and install a production exhaust gas oxygen sensor (one-wire EGO) in the original location in the exhaust manifold.

6.3.6.3  Mount an industrial cooling blower with a nominal air flow rating within 10 000 L/min to 14 000 L/min to blow air vertically over the cast iron exhaust manifold and the manifold exhaust gas oxygen (EGO) sensor. This cooling air is essential to proper EGO operation. Ensure this cooling air is not directed to the engine oil pan or rocker arm cover. Use a deflector shield to prevent air currents at the oil pan. See Annex AX for Safety Precautions.

6.3.6.4  Use the production exhaust pipe front length (minimum 500 mm), including tube collector with shield, leading from the manifold. Route the exhaust from the test cell using accepted laboratory practices. Install an exhaust pressure control valve at any point after the production exhaust pipe to enable the exhaust to be controlled to an absolute pressure. Use of a catalytic converter, or exhaust attenuator, or pipe cooling is optional, provided these devices are installed after the production exhaust pipe front length and specified absolute pressure is maintained. Remove the unused exhaust pipe production fitting, and weld a plate over the opening (see Annex AX).

6.3.6.5  Because this test method is continuously operated at low engine speeds and torque, the water vapor in the exhaust gas tends to condense in the exhaust piping. Install a low point drain in the exhaust piping to remove accumulated water before the start of each test. Depending on the exhaust piping arrangement, if exhaust pressure fluctuations are observed, remove water periodically throughout the 100 h test.

6.3.6.6  Air-To-Fuel-Ratio Sensor—Install a Universal Exhaust Gas Oxygen (UEGO) sensor in the production exhaust pipe to monitor the air-to-fuel ratio. Make a port (30 ± 10) mm downstream of the collector. Orient the UEGO to the front side of the exhaust pipe using the appropriate weld fitting. It is not necessary to direct cooling air over the UEGO sensor.

6.3.6.7  Exhaust Gas Temperature—Measure the exhaust gas temperature using a 6 mm diameter thermocouple. Install the thermocouple in a welded fitting attached to the exhaust pipe at a location (50 ± 10) mm downstream from the end of the collector. Insert the sensor tip to the center of the exhaust pipe (see Annex AX).

6.3.6.8  Exhaust Absolute Pressure—Attach the exhaust pressure sensor tube to a welded fitting installed on the exhaust pipe at a location (50 ± 10) mm downstream from the end of the tube collector. Orient this fitting circumferentially (60 to 90)° from the exhaust temperature sensor.

6.3.6.9  Exhaust Sample Probe—It is optional to install an exhaust sampling probe for emission analyses (percent O2, CO2, CO, HC). If used, locate the exhaust sampling probe 100 mm downstream from the end of the collector on the exhaust pipe. Extend the probe into the center of the exhaust pipe, with the tip of the probe cut to a 45° angle (longest portion facing upstream).

6.3.7  Air-to-Fuel Ratio Control—Control the air-to-fuel ratio (AFR) at a stoichiometric mixture (14.4 ± 0.3) by the engine ECM, using feedback from the production exhaust gas oxygen sensor installed in the exhaust manifold.

6.3.7.1  AFR Measurement—To monitor the reliability of the AFR control, use an AFR analyzer with a separate wide range-sensing element (UEGO) sensor to compute the AFR. Use a Horiba model MEXA 110 lambda analyzer,16 ,12 or the ETAS Lambda Meter LA3.17 ,12 These analyzers are configured to read directly the air-to-fuel ratio. Program the Mexa 110 AFR analyzer with the information shown in Table 4 for the Haltermann KA24E Green test fuel. Input the Mexa 110 analyzer with sensor calibration documentation received with the sensor. It is recommended that a periodic verification of the calibration be performed by exposing the sensor to a 4.0 % O2, N2 balance certified gas. Follow the manufacturer’s calibration procedures for the AFR analyzer used.

TABLE 4 AFR Analyzer ParametersA

| Fuel Properties | Value |
| --- | --- |
| Hydrogen to Carbon ration of the fuel | 1.800 |
| Oxygen Content | 0.000 |

A Stochiometric air-to-fuel ratio for the test fuel is 14.4 to 1.

6.3.8  Ignition System—Do not modify the ignition system for this test method.

6.3.8.1  Monitoring Ignition Timing—Use an automotive timing light (strobe) to visually check the ignition timing.

6.3.9  Engine Coolant System—A schematic diagram of the external coolant system is shown in Annex AX. Use a 50 % deionized water and antifreeze solution, using an extended life ethylene glycol based engine coolant. Texaco Havoline Dex-Cool18 ,12 has been found to meet this requirement (see Annex AX). Configure the plumbing such that the total coolant system capacity, including engine and normal reservoir capacity, is (25 to 30) L. Regulate the system pressure by a 100 kPa radiator-type pressure cap onto the reservoir tank. Plumb the coolant to enter the engine at the thermostat housing (remove the thermostat). Coolant exits the engine at the front of the intake manifold. Circulate a portion of the engine coolant through the specially manufactured jacketed rocker cover (see Annex AX).

6.3.9.1  External Coolant Pump—Use an electric motor-driven centrifugal bronze body pump with a nominal minimum rating of 150 L/min at 100 kPa head pressure. The actual flow range during the test (including break-in) is (20 to 70) L/min.

6.3.9.2  Coolant Heater—Use a nominal 8 kW electric heater, or equivalent external heating source, in the coolant system. This allows engine coolant soak temperatures to be maintained while the engine is not running. Because the ECM coolant temperature sensing system is non-operable, smooth running of the engine upon start-up depends on maintaining the coolant soak temperature.

6.3.9.3  Coolant Heat Exchanger—Use a conventional shell-and-tube heat exchanger for cooling. Flow the engine coolant through the tube side, and use process water on the shell side. A nominal 150 mm diameter by 1200 mm long exchanger has been found to be suitable. Position the heat exchanger vertically, and the coolant inlet at the top of the exchanger. Plumb the high point bleed to remove system air during initial circulation of coolant. Install a sight-glass in the coolant line upstream of the external coolant pump. Plumb a low point drain to allow complete coolant removal.

6.3.9.4  Coolant Control—For control of the coolant out temperature, install an automatic control valve in the process water outlet of the heat exchanger. Use a control valve with a Cv rating of 1.25 for the recommended heat exchanger size.

6.3.9.5  Coolant Flow Control—Measure the coolant flow using a volumetric flow sensor installed in the coolant line between the heat exchanger and the coolant inlet to the engine. A Barco venturi19 ,12 metering element is recommended. Control the flow by an automatic flow control valve on the discharge side of the external pump. A control valve with a Cv rating of 16 is recommended.

6.3.9.6  Jacketed Rocker Cover Coolant System—Route a portion of total coolant system flow through the jacketed rocker cover. Install a tee fitting at the exit of the coolant heat exchanger to allow the coolant flow to split into two circuits (main circuit to the engine thermostat housing and secondary circuit to the jacketed rocker cover (see Fig. 2). The secondary circuit enters the front of the jacketed cover and exits the rear of the cover. Install an automatic air bleed vent near the front of the rocker cover. Limit the secondary circuit flow rate at the exit by installing a two-way control valve, 13 mm nominal internal diameter size, with a flow coefficient rating (Cv) of 1.25. Configure the control valve in the fail-safe open position. The secondary flow joins the primary flow at the suction of the coolant system-circulating pump. Refer to the schematic of the cooling system located in Annex AX.

FIG. 2  Jacketed Rocker Cover

6.3.10  Crankcase Ventilation System (Fig. 3)—Alter the Nissan production routing of the crankcase gasses to ensure that a certain mass flow rate of fresh air is supplied to the valve-train underneath the jacketed rocker cover. Take humidity-conditioned air from the bottom, left rear of the air cleaner housing and route to the rear right side of the rocker arm cover and to the engine front cover.

FIG. 3  Typical Crankcase Ventilation

6.3.10.1  Draw the crankcase off-gas from the engine at the production breather and oil separator. From the breather, the crankcase gas flows through the Positive Crankcase Ventilation (PCV) valve to the bottom plenum of the intake manifold (see Annex AX) for a drawing of the ventilation system plumbing.

6.3.10.2  Use a mass flow meter to measure the fresh airflow to the rocker cover of 10.0 L/min (SLPM, Standard Litres per Minute). This meter, corrected to standard conditions, shall have an accuracy of ± 0.25 L/min (SLPM) at 10 L/min (SLPM). Full scale of the meter shall be a minimum of 20 L/min (SLPM). Time response of the measurement shall be less than or equal to 1.0 s. One model that meets these specifications is Sierra Mass Flow Meter, model 730-N2-1E0PV1V4 (air; 20 SLPM).20 ,12

6.3.10.3  Prior to the meter, install a three-way control valve having a nominal size of 13 mm and a flow coefficient rating of 2.5 Cv. Configure the valve so that loss of control power routes all air to the rocker cover. A Badger Meter 1/2  in. research valve with Trim A meets these requirements.21 ,12 Use a 20 L nominal surge at the exit of the flow meter.

6.3.10.4  The plumbing from the 3-way valve to the engine front cover is a nominal diameter of 10 mm; see Fig. 4. The plumbing from the 3-way valve, through the flow meter and surge chamber, and on to the rear of the rocker cover, is a nominal diameter of 16 mm. A 3.2 mm needle-valve may be installed between the intake and the PCV.

FIG. 4  Location of Crankcase Ventilation Port to Engine Front Cover

6.3.10.5  Diversion for Blowby Measurement—To facilitate the periodic measurement of engine blowby, install a 3-way valve in the hose between the engine PCV and the intake manifold vacuum source. Use a longer hose to connect the rocker cover to the air cleaner housing. During blowby measurement, position the 3-way valve and hoses to route blowby from the rocker cover (bypassing the air cleaner), through the blowby meter, through the 3-way valve, then to the intake manifold vacuum source. Monitor crankcase pressure at the dipstick tube. During blowby measurement, adjust the blowby measurement apparatus for zero crankcase pressure.

6.3.11  Temperature Measurement—Temperature measurement equipment and locations for the required temperatures are specified in 6.3.11.1 – 6.3.11.7. The TMC shall approve alternative temperature measurement equipment. The accuracy and resolution of the temperature measurement sensors and the complete temperature measurement system shall follow the guidelines in Specification E230.

6.3.11.1  Thermocouples—All thermocouples except the intake-air thermocouple shall be premium, sheathed, types with premium wire. The intake-air and ambient air thermocouples may be an open-tip type. Grounded thermocouples may provide a more accurate reading, in situ, when immersion depths are limited. Using grounded thermocouples requires the incorporation of signal conditioning modules for providing electrically isolated inputs to digital computer systems. Use thermocouples of 3.2 mm or 6.4 mm diameter in specific locations. The 3.2 mm thermocouples are specified at locations, which require short immersion depths to prevent undesirable temperature gradients. For exhaust gas temperature, the 6.4 mm diameter thermocouple is recommended. Match thermocouples, wires, and extension wires to perform in accordance with the special limits of error as defined by Specification E230. Either Type J (Iron-Constantan) or Type T (Copper-Constantan) or Type K (Chromel-Alumel) thermocouples are acceptable; Type J is preferred.

6.3.11.2  Resistance Thermometer Detectors—Do not use Resistance Thermometer Detectors (RTDs) as alternatives to thermocouples, due to inherent signal attenuation characteristics that are different from sheathed, grounded thermocouples.

6.3.11.3  Engine Coolant Inlet—Install the engine coolant inlet temperature sensor at the inlet pipe, (200 ± 20) mm from the end of the thermostat-housing nipple. Locate the sensor tip at the center of the pipe inner diameter.

6.3.11.4  Engine Coolant Outlet—Install the engine coolant outlet temperature sensor at the coolant water outlet passage at the front end of the intake manifold. Locate the existing port at the top of the manifold, (50 ± 10) mm from the intake gasket surface. Locate the sensor tip in the center of flow. The recommended thermocouple diameter is 3.2 mm. This temperature is the coolant control point.

6.3.11.5  Engine Oil Gallery Temperature—Precisely weld a thermocouple fitting to the oil filter block (see Fig. 5). A 3.2 mm diameter thermocouple, or equivalent is recommended. Position the sensor tip in the center of the oil passageway. Do not use the engine oil gallery temperature for oil temperature control.

FIG. 5  Oil Filter Block

6.3.11.6  Engine Oil Sump Temperature—Sense the engine oil sump temperature by modifying the drain plug location of the oil pan for a thermocouple fitting, as shown in Fig. 6. Insert the sensor tip (50 ± 5) mm inside the interior surface of the oil pan. Only monitor this temperature. It is not used for oil temperature control.

FIG. 6  Oil Pan

6.3.11.7  Cylinder Head Oil Temperature—Assess the cylinder head oil gallery from the intake side of the head through the vertical passage, centered front-to-rear (see Fig. 7). Drill an access port in a bossed area of the head, and locate 10 mm upward from the deck surface of the head. Drill and tap the access port to accept a 3.2 mm close pipe nipple. Connect a 3.2 mm pipe tee to the nipple. Use a straight-through tee connection for the temperature sensor. Insert the sensing tip into the center of the oil gallery passage. Orient the right angle tee connection downward and use for the measurement of cylinder head gallery pressure. The cylinder head oil temperature is the primary control point.

FIG. 7  Cylinder Head

6.3.11.8  Dynamometer Load Cell Temperature—Measure the dynamometer torque using a strain-gage transducer attached to the moment arm; it is recommended that the environment surrounding the transducer be maintained at a constant temperature. Strain-gage transducers are very sensitive to temperature changes. Use a temperature sensor located near the transducer to monitor ambient variations.

6.3.11.9  Rocker Cover Inlet and Outlet Temperature—Measure the rocker cover inlet and outlet temperature within 150 mm of the inlet and outlet connections. Install the sensor tip in the center of flow.

6.3.11.10  Rocker Cover Gas Temperature—Insert the rocker cover gas temperature sensor through the rear cylinder head rubber gasket (half-moon rubber plug). Drill a 2 mm diameter hole in the rear rubber plug, 4 mm down from the top, flat surface, centered horizontally. Press fit a 3.2 mm diameter closed tip type J thermocouple, 4 cm length, into the drilled hole so that the tip of the sensor is 12 mm from the inside surface of the rubber plug.

6.3.12  Pressure Measurement Equipment—The seven required pressure measurement parameters are shown in 6.3.11.2 – 6.3.11.7. This test method does not specify specific measurement equipment. This allows reasonable opportunity for adaptation of existing test stand instrumentation. The accuracy and resolution of the pressure measurement sensors and the complete pressure measurement system shall follow the guidelines detailed in ASTM research report RR:D02-1218.22

6.3.12.1  Allowance for Pressure Head Deviations—Use tubing between the pressure tap locations and the final pressure sensors and incorporate condensate traps. This is important in applications where low air pressures are transmitted through lines passing through low-lying trenches between the test stand and the instrument console. Locate the pressure transducer at the same elevation as the measurement location, or account for the pressure head. Design the oil pressure sensor / tubing lines to minimize the trapped oil volume in the tubing lines.

6.3.12.2  Intake Manifold Vacuum—Measure the intake manifold vacuum on the throttle body, at an existing port located just below the throttle plate.

6.3.12.3  Engine Oil Pressure—Sense the engine oil pressure at the production location on the oil filter block (see Fig. 5). Route the sensing line to a tee fitting, allowing ports to a pressure transducer and an analog pressure gage.

6.3.12.4  Cylinder Head Oil Gallery Pressure—Assess the cylinder head oil gallery from the intake side of the head through the vertical passage, centered front-to-rear. Drill this access port in a bossed area of the head that is located 10 mm upward from the head deck surface. Drill and tap the access port to accept a 3.2 mm close pipe nipple. Connect a 3.2 mm pipe tee into the nipple. Orient and use the right angle tee connection downward for the measurement of cylinder head gallery pressure.

6.3.12.5  Fuel Pressure—Measure the fuel pressure through a cross-fitting port located in the line between the fuel pump and the fuel rail inlet.

6.3.12.6  Crankcase Pressure—Attach the crankcase pressure sensing line to a fitting welded to the dipstick tube. Locate this fitting approximately (80 ± 10) mm from the top of the dipstick tube (see Fig. 8). The sensor shall be capable of measuring positive and negative pressure. If using a manometer, install a liquid trap to prevent manometer fluid from entering the crankcase.

FIG. 8  Dipstick Tube

6.3.12.7  Coolant Pressure—Attach the coolant pressure sensing line to a fitting welded to the coolant reservoir.

6.3.13  Flow Rate Measurement Equipment—Measure the engine coolant, fuel, and blowby flow rate. The accuracy and resolution of the flow rate measurement sensors and the complete flow rate measurement system shall follow the ASTM research report RR:D02-1218.22

6.3.13.1  Engine Coolant Flow Rate—Determine the engine coolant flow rate by measuring the differential pressure drop across a venturi flowmeter. A Barco #70523 is suitable. Calibrate a differential pressure transducer to provide an output (litres per minute). Take precautions to prevent air pockets from forming in the lines to the pressure sensor. Transparent lines or bleed lines, or both, are beneficial in this application.

6.3.13.2  Fuel Consumption Rate—Determine the fuel consumption rate using a mass flow meter installed in the makeup fuel line. The flow meter output shall allow real-time measurement of fuel rate (kilograms per hour) and total fuel consumed (kilograms). A MicroMotion Model D-6 is satisfactory.

6.3.13.3  Blowby Flow Rate Measurement System—Use the apparatus shown in Annex AX for measurement of the blowby flow rate. The measurement system routes the blowby through an external, sharp-edge orifice and into the engine intake manifold by way of an auxiliary PCV valve. Maintain the crankcase gage pressure at (0.0 to 0.025) kPa during system operation to minimize crankcase leakage. Determine the blowby flow rate by measuring the differential pressure drop across the sharp-edge orifice. Use an inclined manometer or differential pressure sensor for the orifice differential pressure measurement. The crankcase pressure sensor shall have a (0 to 1) kPa range and be adequately damped to indicate a zero gage pressure. The sharp-edge orifice is specifically designed for blowby flow rate measurement and shall be fabricated in strict compliance with the specifications available from the TMC. The assembly contains five orifices. Use a 3.175 mm orifice for the blowby flow rate range of (5 to 12 L/min). The flow rate measurement orifice fabrication location and the blowby flow rate measurement can be obtained from the TMC.

6.3.14  Process Cooling Water—Provide the engine jacket coolant heat exchanger, the oil cooler, and the eddy-current dynamometer with process cooling water to maintain proper operating temperatures.

6.3.14.1  Dynamometer Cooling System—Water-cool the eddy current electric dynamometer. Provide provisions for automatic test shutdown in the event the dynamometer cooling water is shut off.

6.3.15  Volumetric Graduates—Use volume to measure the test oil quantity. Do not use mass, converted to a calculated volume. Recommend using a 1000 mL graduate and 500 mL graduate for volumetric determinations. All volumetric graduates shall be accurate to 2 % of full scale. The large graduate resolution shall be 100 mL, and the small graduate resolution shall be 25 mL.

6.4  Test Engine Hardware—This section specifies the hardware required to build the test engine. Conduct the engine break-in procedure prior to the first test on a long block and on the first test when a cylinder head is replaced. The new engine is a long-block, as received. Use the camshaft and rocker arms in the new engine for break-in purposes only. Remove and modify the new cylinder head for the cylinder head oil gallery temperature and pressure measurement port, and for valve spring force calibration. Clean and reassemble the head using the break-in camshaft and rocker arms. Use the break-in procedure shown in 11.1.3. After break-in, replace the break-in camshaft and rocker arms with the new camshaft and rocker arms parts.

6.4.1  Nissan Supplied Component Kits—Obtain the test parts and engines for this test method from Nissan North America11 ,12 (see Annex AX).

6.4.1.1  Test Engine Long-Block—Order the test engine long-block assembly (also called bare engine assembly) as shown in Annex AX. The test engine includes the block, pistons, rods, crankshaft, oil pan, front cover, cylinder head, and rocker arm cover final assembly. Use the camshaft and rocker arms during engine break-in only; but they are not official test parts Annex AX contains information regarding the number of test the short-block may be used. Use the original cylinder head for the number of test listed in Annex AX.

6.4.1.2  Stand Set-Up Kit—There are four component kit parts that comprise the stand setup kit (see Annex AX). These four component kits include crankshaft pulleys, flywheel, intake and exhaust manifolds, air cleaner, fuel injection system, EGR block-off plate, ignition distributor, wiring, starter motor, fuel pump, exhaust pipe, and oil cooler. Use one of each of the four component kit parts to configure one test installation.

6.4.1.3  Test Kit—For every official test conducted, use one test kit. This kit includes a camshaft, two rocker shafts, twelve rocker arms, three oil filters, and four spark plugs. The test kit camshaft and rocker arms are considered critical test parts. Do not substitute critical test parts with any other dealer, or aftermarket supplied hardware. Do not alter or modify any test kit part from as-received condition unless expressly approved by the surveillance panel. If any test kit part is modified, note that in the comment section of the test report.

6.4.1.4  Cylinder Head Replacement Kit—Every engine short-block is used for the number of tests listed in Annex AX. Use the original cylinder head for the number of tests listed in Annex AX. After that number of tests, install a replacement cylinder head for the number of runs listed in Annex AX. To assemble and install the bare cylinder head, use 1 gasket and seal kit. Install new calibrated valve springs, intake and exhaust valves with the replacement head (see Annex AX). When the replacement head is installed onto the engine, use the original supplied camshaft and rocker arms for conducting another break-in prior to the next test.

6.4.2  Procurement of Critical Parts—The test camshaft and rocker arms are critical engine parts for this test method. Obtain these parts by annual orders placed through Nissan North America. Do not use dealer, service, or aftermarket camshafts and rocker arms for this test method. Do not alter or modify any critical part from as-received condition unless expressly approved by the surveillance panel. If any critical part is modified, note that in the comment section of the test report.

6.4.3  Required New Engine Parts—This test method is a flush-and-run type test. For each test, the camshaft and rocker arms are replaced. Use all the parts in test kit number 13000-40F85 as shown in Annex AX.

6.4.3.1  Replace the spark plugs for each test, just prior to the oil flush (see 9.7.2). Gap the spark plug at 0.99 mm.

6.4.4  Reusable Engine Parts—Replace the engine short-block and the cylinder head as specified in Annex AX. If the engine demonstrates deterioration (excessive blowby or oil consumption or fuel dilution, poor compression, low oil pressure, clearances beyond service limits, or stripped fasteners) prior to this expected life (see Annex AX), replace the engine and follow the break-in procedure prior to resuming non-reference oil testing. Do not exceed the number of tests on the short-block or cylinder heads listed in Annex AX.

6.4.4.1  Replace the PCV valve, fuel filters, rocker cover gaskets, and air filter element whenever the cylinder head, or engine, or both are replaced. The ignition distributor can be used as long as it remains serviceable. Distributor cap (Nissan Part Number 22162–40F00) and rotor (Nissan Part Number 22157–21E01) can be replaced as needed.

6.4.4.2  Reuse the jacketed rocker arm cover, oil pan, oil cooler, flywheel, intake and exhaust manifolds, throttle body, modified dummy water pump, spark plug wires, fuel injection system components, and engine sensors, as long as they continue to function properly.

6.5  Special Measurement and Assembly Equipment—This section describes the special apparatus, tools, and equipment required for engine measurement and assembly. Routine laboratory and workshop items are not included. Specific engine tools are shown in Table 5.

TABLE 5 Test Tools

Item

3/8-in. Drive Impact Gun

3/8-in. Drive Speed Handle

3/8-in. Drive Rachet

3/8-in. Drive 4-in. Extension

3/8-in. Drive Rachet

Medium Flat Head Screwdriver

Large Flat Head Screwdriver

5/8-in. Wrench, Combination

5/8-in. Spark Plug Socket, 3/8 in. Drive

27 mm Deep Socket, 3/8-in. Drive

24 mm Impact Deep Socket, 3/8-in. Drive

12 mm Deep Socket, 3/8-in. Drive

10 mm Deep Socket, 3/8-in. Drive

Digital Bore Gage with Metric Head

Dial Indicator Set with Magnetic Base

Mounting Plate for Dial Indicator

1-2 in. Digital Micrometer

Spark Plug Gapping Tool

Suction Device (Syringe and Tubing)

Wooden Wedge Tool (see Annex XX)

Utility Rocker Shafts

Pin Vise and 1.17 mm Diameter Drill Bit

6.5.1  Camshaft Lobe Measurement Equipment—Trace the camshaft lobes with a surface texture profilometer system.

6.5.1.1  Use a surface measurement profilometer with real time digital display and graphical output capability. The vertical scale graphical resolution shall be capable of 1 μm per graph division. The profilometer shall be capable of traversing at least 100 mm, with a straightness accuracy equal to or less than 1 mm per 100 mm of traversed length.

6.5.1.2  Use the profilometer pickup without a skid. Use a conical or spherical diamond tip stylus, with a nominal radius of 5 μm.

6.5.1.3  The nominal traversing speed is (0.50 to 0.75) mm/s. A computer interface is recommended. The Precision Devices, Inc. MicroAnalyzer 2000 system,23 ,12 see Fig. 9, a computer-driven profilometer, may be used. Equip the profilometer with custom V-blocks (see Annex AX) for holding the work-piece (the camshaft on its journals), and an optical angle encoder for determining the cam shaft angular position (see Annex AX).

FIG. 9  Cam Lobe Measurement Apparatus

6.5.1.4  View the data from the trace in the profile mode, allowing an analysis of the texture and waviness of the trace. Configure the instrument software for a two-point line texture leveling at the unworn edges of the cam lobe. Use this reference line for wear measurements. Display the profile waviness, using the Gaussian smoothing filter, set at 0.25 mm cutoff length, and do not remove the filter width at the ends of the texture.

6.5.1.5  Base the lobe wear measurement upon the vertical dimension between the horizontally positioned, twp-point leveling line (reference line) and the lowest point in the waviness profile.

6.5.2  Unassembled Valve Spring Calibration Device—Use a device to screen inner and outer valve springs before assembly in the cylinder head. Measure the individual spring forces at various compressed heights according to the Nissan Service Manual.24 ,12 The tester shall be accurate to 2 % and a resolution of 5 N.

7.  Reagents and Materials

Note 6—Use 13 L of the non-reference test oil sample to perform the 100 h Valve-train Wear test.

7.1  Coolant for Engine and Rocker Arm Cover—De-mineralize (less than 0.29 g/kg) the coolant or use distilled water mixed with an extended-life ethylene glycol antifreeze/coolant at a 50:50 ratio.

7.2  Fuel—Use Haltermann KA24E Green test fuel for this test method as shown in Annex AX. (Warning—Flammable. Health hazard.) It is dyed green to preclude unintentional contamination with other test fuels. Use approximately 200 L of fuel for each test (100 cycles). This fuel has a hydrogen-to-carbon ratio of 1.80 to 1.

7.2.1  Fuel Approval Requirements—The fuel is blended as needed by the fuel supplier. Base the fuel batch acceptance upon the physical and chemical specifications given in Annex AX . Engine validation tests are not necessary for fuel batch acceptance.

7.2.2  Fuel Analysis—Monitor the test fuel using good laboratory practices. Analyze each fuel shipment to determine the value of each parameter for existent gum as described in Test Method D381, RVP as described in Test Method D323, and API Gravity as described in Test Method D287. Compare the results to the original values supplied by the fuel supplier. The analytical results shall be within the tolerances shown in parentheses beside each parameter. This provides a method to determine if the fuel batch is contaminated or has aged prematurely. If any analytical result falls outside the tolerances, the laboratory shall contact the fuel supplier for problem resolution.

7.2.2.1  Fuel Deterioration—Analyze the fuel semiannually to ensure the fuel has not deteriorated excessively or been contaminated in storage.

7.2.2.2  Analyze the fuels using Test Methods D287, D323, D381, and D525.

7.2.3  Fuel Shipment and Storage—Ship the fuel in containers with the minimum allowable venting as dictated by all safety and environmental regulations, especially when shipment times are anticipated to be longer than one week. Store the fuel in accordance with all applicable safety and environmental regulations. If the run tank has more than one batch of fuel, document the most recent batch in the test report. Do not top-off the run tank with the new fuel batch unless the tank is less than 10 % full.

7.3  Lubricating Oils:

7.3.1  Break-in Lubricating Oil—An engine break-in procedure (see 11.1.3) is immediately conducted following the replacement of new, major engine components (that is, engine short-block, or cylinder head, or both). Use the proper reference oil, REO 926-2, from the TMC for the break-in procedure. Use approximately 3.5 L of this reference oil for each break-in procedure.

7.3.2  Lubricant for Hydraulic Lash Adjusters—The rocker arms hydraulic lash adjusters may ingest air during shipping and handling. Prior to installing the rocker arms in the test engine, prime the lash adjusters with an SAE 20 API SA grade oil, as shown in SAE J183. Place the rocker arms on their side, for a minimum of 1 h, in a container filled with the SA grade oil to allow trapped air to bleed out. If using a vacuum chamber, reduce the minimum soak time to 5 min. Immediately install the rocker arms in the engine after the rocker arms are removed from the oil-filled container. Do not allow the rocker arms to lie on their side after air has been bled.

7.3.3  Short-block Assembly Lubricant—For engine short-block inspection and reassemble, use SAE 20 API SA grade oil as the assembly lubricant.

7.4  Miscellaneous Materials:

7.4.1  Solvents and Cleansers—No substitutions for 7.4.1.1 – 7.4.1.3 are allowed. Use adequate safety provisions with all solvents and cleaners.

7.4.1.1  Solvent—Use only mineral spirits meeting the requirements of Specification D235, Type II, Class C for Aromatic Content ((0 to 2) vol %), Flash Point (61 °C, min) and Color (not darker than +25 on Saybolt Scale or 25 on Pt-Co Scale). (Warning—Combustible. Health hazard.) Obtain a Certificate of Analysis for each batch of solvent from the supplier.

7.4.1.2  Pentane—(Warning—Flammable. Health hazard.) Available from petroleum solvent suppliers.

7.4.1.3  Cylinder Block and RAC Cleaning Detergent—Tri-sodium phosphate and any commercial coolant cleanser. (Warning— Caustic. Health hazard.)

7.4.2  Sealing Compounds—Use a silicone based gasketing compound during engine assembly (for example, oil pan). Use only the silicone gasket shown in Annex AX .

TABLE 6 Bolt Torques

| Bolt Types | Torques |
| --- | --- |
| Cylinder head bolts | 82° beyond 29 N·m |
| Camshaft Sprocket bolt | 137 N·m |
| Intake Manifold bolts | 19 N·m |
| Camshaft cap bolts | 39 N·m |
| Exhaust Manifold bolts | 22 N·m |
| Rocker Cover bolts | 9 N·m |

8.  Oil Blend Sampling Requirements

8.1  Sample Selection and Inspection—The non-reference oil sample shall be uncontaminated and representative of the lubricant formulation being evaluated.

Note 7—If the test is registered using the American Chemistry Council25 protocols, the assigned oil container formulation number shall match the registration form.

8.2  Non-reference Oil Sample Quantity—Use a minimum of 13 L of new oil to complete the Sequence IVA test, including the oil flushes. Normally the supplier provides a 15 L new oil sample to allow for inadvertent losses.

8.3  Reference Oil Sample Quantity—The TMC provides a 15 L reference oil sample for each stand calibration test.

9.  Preparation of Apparatus

Note 8—This section details those recurring preparations necessary for test operation. This section assumes the engine test stand facilities and other hardware described in Section 6 are in place.

9.1  Test Stand Preparations:

9.1.1  Instrumentation Calibration—Calibrate all sensors and indicators before or during the test for the type instrumentation used. See Section 10 for the calibration requirements.

9.1.2  Oil Cooler Cleaning—Use clean mineral spirits followed by forced-air drying to clean the oil cooler, if required (see 6.1.3).

9.1.3  Air Cleaner Filter—Replace the air cleaner filter element every eight tests, or more frequently if intake air pressure is insufficient.

9.1.4  Draining Exhaust Piping—Prior to the start of each test, drain the low point of the exhaust piping to eliminate water accumulation. Drain water during a test if exhaust pressure control becomes unstable.

9.1.5  External Hose Replacement—Inspect all external hoses used on the test stand and replace any hoses that have become unserviceable. Check for internal wall separations that could cause flow restrictions. Inspect and replace the oil cooler coolant hoses when the oil cooler is replaced (every 16 tests).

9.1.6  Stand Ancillary Equipment—Service the dynamometer and driveline components, as required. The dynamometer torque measurement shall be accurate (no unaccounted forces from hoses, load cell temperature gradients, or trunnion bearing hysteresis).

9.2  General Engine Assembly Preparations—Follow and complete the assembly preparation functions shown in Table 6. Different laboratories and engine assemblers may complete the functions in slightly different sequences. Perform the functions in a specific manner or at a specific time in the assembly process as shown in Table 6. Complete any assembly instructions not detailed below according to the instructions in the Nissan Service Manual. Follow the bolt torque specifications shown in Table 6.

9.3  Cylinder Head Preparations—Modify, clean, assemble, and calibrate new test engine cylinder heads before using the cylinder head for testing purposes. Record the measurements shown in 9.3.3 on the appropriate form.

9.3.1  Cylinder Head Modification—Assess the cylinder head oil gallery from the intake side of the head through the vertical passage, centered front-to-rear. Drill this access port in a bossed area of the head and located (10 ± 1) mm upward from the deck surface of the head. Drill and tap the access port to accept a 3.2 mm close pipe nipple.

9.3.2  Cylinder Head Cleaning—Complete the modification shown in 9.3.1, then thoroughly clean the cylinder head using mineral spirits spray. Clean the valve guide bores and oil passages using a nylon bristle brush in conjunction with the mineral spirits spray (see 6.1.3). Then rinse the head with a clear water spray, and dry with forced air.

9.3.3  Camshaft Journal Clearance—Use only the vertical measurements to determine the overall cam journal to bearing clearance (front and rear). This clearance may not exceed 0.120 mm. Replace the cylinder head or camshaft, or both, if the clearance exceeds 0.120 mm.

9.3.3.1  Camshaft Bearing Bore Measurements—Install cam-bearing caps, without the camshaft, with rocker shafts and tighten the lubricated cap bolts to 39 N·m.

(1)  When the cylinder head is new, perform a comprehensive set of measurements of the five cam bearing bores to ensure the bores are sized to specification, and that the bores are round and not tapered. Measure the front and rear of each cam bore in three directions: vertical and two measurements at 45° from vertical. This results in six measurements for each cam bore. The standard inner diameter specification is (33.000 to 33.025) mm. Run-out and taper shall not exceed 0.025 mm.

(2)  After the first test on a cylinder head, use only the vertical dimension of the front and rear of each cam bore (from the comprehensive measurements performed on the new cylinder head) to determine the cam journal to bearing bore clearance before subsequent tests.

9.3.3.2  Camshaft Journal Measurements—Measure the front and rear outer diameters of each of the five camshaft journals. The standard outer diameter is (32.935 to 32.955) mm. Use these measurements to calculate cam journal to bearing bore clearance.

9.3.3.3  Camshaft End Play—With the camshaft installed in the cylinder head, measure the camshaft endplay with a dial indicator. If the endplay exceeds (0.07 to 0.20) mm, do not use that cylinder head/camshaft combination.

9.3.4  Cylinder Head Assembly—Assemble the cylinder head using the instructions in the Nissan Service Manual. Lap the valve faces to their respective valve seats to ensure a proper seal. Using a new cylinder head precludes replacing or machining the valves, valve guides, and valve seats. Use SAE 20 API SA lubricant to lubricate the valve stems and valve guides upon assembly.

9.3.5  Initial Valve Spring Screening—Measure the force on the unassembled valve spring calibration device. The valve spring parameters shall be within the specifications show in Table 7. Determine the final valve spring assembled force with the valve springs installed in the cylinder head (see 9.3.6). Springs slightly outside the unassembled force specification can be within the assembled force specification when installed in the cylinder head using shims under the springs.

TABLE 7 Valve Spring Specifications

| Parameters | Intake | Exhaust |
| --- | --- | --- |
| Force |  |  |
| outer | 604.1 N at 37.6 mm | 640.4 N at 34.1 mm |
| inner | 284.4 N at 32.6 mm | 328.5 N at 29.1 mm |

9.3.6  Installed Valve Spring Calibration—Lubricate each valve seal and valve stem with SAE 20 API SA oil. Install the valve seal over the end of the valve stem with a plastic installation cap in place. Carefully seat the seals fully on the guides. Install pre-screened valve springs and retainers. When installing the valve springs and retainers, do not compress the springs excessively. Excessive spring compression may damage the valve seals.

9.3.6.1  Assembled Force Calibration—Measure and record the assembled valve spring loading with the valve springhead calibration fixture shown in Fig. 10 and the part number in Annex AX.

FIG. 10  Head Calibration

9.3.6.2  The procedure detailed below includes measuring each installed valve spring at two deflection points (see Table 8).

TABLE 8 Spring Specifications

| Intake Valve Spring | Exhaust Valve Spring |
| --- | --- |
| Force | Deflection | Force | Deflection |
| (889 ± 35) N | 9.86 mm | (969 ± 35) N | 9.86 mm |
| (438 ± 35) N | 1.27 mm | (447 ± 35) N | 12.7 mm |

(1)  Check the apparatus load cell calibration for accuracy.

(2)  Install the cylinder head holding fixtures (see Annex AX) for the test cylinder head.

(3)  Position the test cylinder head in the holding fixture with intake valve springs accessible.

(4)  Install the rocker cover gasket rail extension bracket on the test cylinder head.

(5)  Starting with the intake valve spring No. 1 (far right), position the air cylinder and load cell on valve tip to allow compressing of the valve spring.

(6)  Position the dial indicator with the plunger on the rocker cover gasket rail extension bracket.

(7)  With no air pressure to the cylinder, the dial indicator shall read between (2.5 and 5.00) mm deflection pre-load against the rocker cover gasket rail extension bracket. This will allow determination of positive or negative valve displacement.

(8)  Adjust the dial indicator face to position the needle at the zero mark with no air pressure to the air cylinder.

(9)  Adjust the air regulator to provide enough air pressure to the air cylinder for compression of the valve spring to occur. If a valve spring has already been measured, then the air regulator should already be adjusted.

(10)  Actuate and discharge the air cylinder in a rapid and consistent manner, several times, to compress the valve spring and assure proper movement of the valve-train components.

(11)  Actuate the air cylinder and adjust the air regulator to decrease the air pressure to allow the valve to close to the fully seated position.

(12)  Adjust the air regulator to gradually apply force to achieve exactly 1.27 mm valve opening. Record the indicated force in newtons at 1.27 mm.

(13)  Continue to adjust the air regulator to gradually apply force to achieve exactly 9.86 mm valve opening. Record the indicated force in newtons at 9.86 mm.

(14)  Discharge the air cylinder without adjusting the air regulator, and allow the intake valve to close to the fully seated position. Position the air cylinder/load cell to check the remaining seven-intake valve springs. Conduct Steps 5–13 for each remaining intake valve spring.

(15)  Position the test cylinder head in the holding fixture with the exhaust valves accessible.

(16)  Install the rocker cover gasket rail extension bracket on the cylinder head, and position the air cylinder and load cell to check the four exhaust valve springs. Conduct Steps 5–13 for each exhaust valve spring.

(17)  Replace any springs that are too strong. Use shims to adjust springs to a higher installed force. Recheck spring calibration and record indicated force, in newtons, of any replaced or shimmed springs.

9.3.7  Vacuum Check—Prior to installing the cylinder head on the engine, conduct a vacuum check to ensure no leakage past the valves, and so forth. The method of checking is left up to the laboratory.

9.4  Short Block Preparations—Not normally required.

9.5  Short Block Assembly—Use SAE 20 API SA grade oil as the assembly lubricant for inspection and reassemble the engine short-block.

9.6  Final Engine Assembly—Use the engine short-block assembly as received from Nissan. Alternatively, disassemble, inspect, and reassemble the engine short-block according to the Nissan Service Manual. For this test, maintain the piston ring end gaps to factory specifications. Do not increase the ring end gaps to elevate engine blowby. Observe the corrected blowby rate at (6.0 to 10.0) L/min during the Stage I operating condition.

9.6.1  Cylinder Head Installation—The parts kit (see Annex AX) include the assembled new cylinder head. The valve seats and valve faces are machined to Nissan production specifications. Laboratory inspection may reveal the need to lap the valve seat and valve faces. Perform the following before assembling the cylinder head:

9.6.1.1  Stamp with a laboratory specific identification code.

9.6.1.2  Modify for measurement access to the cylinder head oil gallery (see 6.3.11.7).

9.6.1.3  Measure for cam bearing bore diameter, specification (33.000 to 33.025) mm.

9.6.1.4  Clean with mineral spirits and forced-air dried (see 6.1.3).

9.6.1.5  Reassemble according to the Nissan Service Manual. Use SAE 20 API SA grade oil to lubricate valve guides and valve stems. Install new valve seals.

9.6.1.6  Measure and adjust installed valve spring force to test specifications (see 9.3.5). Install the modified and measured cylinder head on the short-block using a new cylinder head gasket (see Annex AX). Torque the head bolts in proper sequence, and in five stages according to the Nissan Service Manual. Use an angle-meter torque wrench to properly torque the cylinder head bolts.

9.6.2  Pre-Test Cam Shaft Inspection and Measurement:

9.6.2.1  Visually inspect the pre-test camshaft lobes and bearing journals for rust, scratches, gouges, chipped areas, and other surface defects. Reject camshafts that are judged unsuitable for test.

9.6.2.2  Check the pre-test camshaft for straightness. With the camshaft supported by V-blocks at journal No(s). 1 and 5, measure the camshaft run-out at the center journal. Reject the camshaft if the total indicator reading of the run-out exceeds 20 μm.

9.6.2.3  Using a profilometer to perform pre-test measurements, check the camshaft for lobe concavity, convexity, and taper across the nose of each camshaft lobe. Reject any camshaft that exhibits concavity or convexity of more than 4 μm in height or depth or taper of more than 10 μm in variation. The laboratory has the one-time option of sending any camshaft, including camshafts rejected per the above criteria, at their discretion, to OHT for regrinding. OHT regrinds camshafts according to Nissan specifications. All camshafts meeting the above specification can then be put into service. Any test run with a reground camshaft should be noted in the test report comment section.

9.6.3  Camshaft Installation—After modifying, cleaning, assembling and calibrating the cylinder head, install the measured test camshaft. Clean the camshaft with mineral spirits and forced-air dry (see 6.1.3). Coat the journals with SAE 20 API SA grade oil or new test oil. Coat the lobes with new test oil only. Do not use REO 926-2 or SAE 20 API SA to lubricate the lobes.

9.6.3.1  Install the test camshaft, camshaft journal caps and dummy rocker shafts, and torque to 39 N·m. Install the camshaft sprocket and snug the sprocket bolt (do not torque to specification).

9.6.3.2  Measure and record the camshaft endplay and the camshaft sprocket run-out (see Annex AX).

9.6.3.3  Temporarily remove the test camshaft, camshaft sprocket, camshaft journal caps, and utility rocker shafts.

9.6.3.4  Pre-fill the cavities of the cylinder head under the camshaft with new test oil.

9.6.3.5  Oil the camshaft journals and journal bores with SAE 20 API SA or new test oil and install the test camshaft.

9.6.3.6  Install the camshaft journal caps in their proper positions.

9.6.3.7  Install the rocker shaft assemblies and loosely screw the bolts. Do not tighten the bolts.

9.6.3.8  Install the camshaft sprocket and timing chain, aligning the punch-mark with the glaring link. Torque the camshaft bolt to 135 N·m using the appropriate cam sprocket holding tool.

9.6.3.9  Center each individual rocker arm on its respective cam lobe, and then snug the ten rocker shaft bolts and torque to 39 N·m using the torque sequence in the Nissan Service Manual.

9.6.3.10  After installing the camshaft pour new test oil over the rocker arms, rocker shafts, and camshaft. Excess oil will drain through the open oil pan drain-valve.

9.7  Replacement of Valve-train Wear Parts—The valve-train wear tests is a flush-and-run type procedure, except replace the valve-train test parts for each test. Use the following procedure when replacing these parts:

9.7.1  Removal of Valve-train Wear Parts:

9.7.1.1  Disconnect wiring or hoses that hinder access to the rocker cover.

9.7.1.2  Remove the crankshaft pulley guard.

9.7.1.3  Remove the blowby fresh airline and the spark plug wire loom from the rocker cover.

Note 9—At Southwest Research Institute (SwRI), client test oil is poured over valve-train parts at test start before Flush 1, NOT during parts installation.

9.7.1.4  Remove the rocker cover.

9.7.1.5  Remove all four spark plugs.

9.7.1.6  Rotate the engine by hand to set piston No. 1 at top dead center on its compression stroke (align the timing indicator with 0° at the yellow mark on the crankshaft pulley). Position the dowel pin on camshaft front at the 12:00 position and align the punch-mark on the camshaft sprocket with the glaring link on the timing chain. This may require rotating the engine several times until the mark and glaring link line up.

9.7.1.7  Remove the front hoist bracket from the cylinder head. It may hang from the ground straps.

9.7.1.8  Remove the half-moon rubber plug from the front of the cylinder head.

9.7.1.9  Install the timing chain wooden wedge tool between timing chain at the tensioner using a large screwdriver to set in place.

9.7.1.10  Remove the camshaft bolt and sprocket using a 24 mm impact deep socket.

9.7.1.11  Remove the ten rocker shaft bolts according to the Nissan Service Manual.

9.7.1.12  Carefully remove the rocker shafts and rocker arms. Avoid touching the rocker arm pads. The rocker arms shall remain assembled on the rocker shafts.

9.7.1.13  Remove the five camshaft journal caps. Number the caps for their location.

9.7.1.14  Carefully remove the test camshaft without damaging the lobes. Avoid touching the camshaft lobes.

9.7.1.15  Using a suction device, remove the used test oil that is trapped in the cavities of the cylinder head under the camshaft. Do not add this oil to the drained test oil. Properly discard this oil.

9.7.1.16  Bring the end-of-test parts, including the assembled rocker shaft assemblies to the engine disassembly area. Before disassembling the used rocker shaft assemblies, properly label the rocker arms.

9.7.1.17  Disassemble the rocker shaft assemblies.

9.7.1.18  Remove the reusable retainer clips from the used rocker shafts.

9.7.1.19  Clean the end-of-test camshaft and rocker arms with mineral spirits and pentane if necessary, then forced-air dry (see 6.1.3). Deliver the cleaned camshaft to the measurement room, and the cleaned rocker arms to the rating area or place them in a desiccator.

9.7.2  Installing New Valve-train Test Parts:

9.7.2.1  Obtain the new test parts, including the pre-measured test camshaft. See Annex AX for a new parts listing with their corresponding part numbers for each turnaround.

9.7.2.2  Number the new rocker arms E1 to E4 and I1 to I8.

9.7.2.3  Clean the new camshaft, rocker shafts and rocker arms, as well as the used retainer clips, camshaft journal caps, camshaft sprocket, rocker cover and all bolts with mineral spirits and forced-air dry (see 6.1.3).

9.7.2.4  Perform and record camshaft bearing journal diameter measurements.

9.7.2.5  Perform a bleed down of the rocker arms with built in hydraulic lash adjusters. Place the rocker arms in a container or vacuum chamber, fully submerging them in SAE 20 API SA grade oil. Using a vacuum chamber reduces the minimum soak time from 1 h to 5 min. Once the rocker arms are air bled, do not place them on their sides. Orient the rocker arms in their natural engine position.

9.7.2.6  Assemble the new rocker arms and used retainer clips on the new rocker shafts according to the Nissan Service Manual.

9.7.2.7  Install the five camshaft journal caps and bare rocker shafts, using the designated utility rocker shafts, and torque to 39 N·m .

9.7.2.8  Perform and record camshaft bearing bore diameter measurements.

9.7.2.9  Remove the camshaft journal caps and utility rocker shafts.

9.7.2.10  Install the test camshaft, camshaft journal caps and dummy rocker shafts, and torque to 39 N·m. Install the camshaft sprocket and snug the sprocket bolt, and do not torque to specification.

9.7.2.11  Measure and record the camshaft endplay and the camshaft sprocket run-out according to the Nissan Service Manual.

9.7.2.12  Remove the test camshaft, camshaft sprocket, camshaft journal caps, and dummy rocker shafts.

9.7.2.13  Pre-fill the cavities of the cylinder head under the camshaft with new test oil.

9.7.2.14  Oil the camshaft and journal bores with new test oil and install the test camshaft.

9.7.2.15  Install the camshaft journal caps in their proper positions.

9.7.2.16  Install the rocker shaft assemblies and loosely screw the bolts. Do not yet tighten the bolts.

9.7.2.17  Install the camshaft sprocket and timing chain, aligning the punch-mark with the glaring link. Using the appropriate cam sprocket holding tool, torque the camshaft bolt to 135 N·m.

9.7.2.18  Center each individual rocker arm on its respective cam lobe, then snug the ten rocker shaft bolts and torque to 39 N·m using the torque sequence as shown in the Nissan Service Manual.

9.7.2.19  After installation pour new test oil over the rocker arms, rocker shafts and camshaft. Excess oil will drain through the open oil pan drain-valve.

9.7.2.20  Inspect the half-moon rubber plugs found on the front and rear of the cylinder head. Replace the plugs if necessary.

9.7.2.21  Install the half-moon rubber plug on the front of the cylinder head. Apply the proper amount of the Nissan silicone gasket maker to the bottom of the plug. Repeat for the half-moon rubber plug found on the rear of the cylinder head.

9.7.2.22  Inspect the rocker cover gasket, and replace if necessary.

9.7.2.23  Install the rocker cover and torque in sequence to 8 N·m as shown in the Nissan Service Manual.

9.7.2.24  Reinstall the blowby fresh airline on to the rocker cover and tighten the clamp.

9.7.2.25  Install spark plug wire loom and reconnect any other wiring or hoses.

9.7.2.26  Install the front hoist bracket and torque the nuts to 20 N·m.

9.7.2.27  Install new spark plugs (see Annex AX). Gap the plugs to 0.99 mm and torque to 14 N·m.

9.7.2.28  Install the front pulley guard.

10.  Data Acquisition, Reference Oil Application, and Equipment Calibration and Maintenance

10.1  Data Acquisition:

10.1.1  Computer Data Acquisition—The test stand should log operational data using a computer data acquisition system with sensor configurations process is described in 10.1.2 – 10.1.4.

10.1.2  Frequency of Logged Test Cycle Data—Log the test cycle data at a sampling rate of 1-Hz.

10.1.3  Signal Conditioning—Do not exceed the controlled operational parameters for system time response as shown in Table 2. The system time response includes the total system of sensor, transducer, analog signal attenuation, and computer digital filtering. Use single-pole type filters for attenuation. For temperature sensors only grounded thermocouples are acceptable.

10.1.3.1  Isolated Inputs—Use signal-conditioning modules to provide isolated inputs to the digital computer.

10.2  Reference Oil Application:

Note 9—10.2.6 and 10.2.7 and Annex A1 - 4 describe the involvement of the TMC in respect to calibration procedures and acceptance criteria for a testing laboratory and a test stand, and the issuance of Information Letters and memoranda affecting the test method.

10.2.1  Testing of Reference Oils—Periodically conduct tests on reference oils according to the following:

10.2.1.1  Conduct reference oil tests on each calibrated test stand within a laboratory according to TMC guidelines.

10.2.1.2  Obtain reference oils directly from the TMC. These oils are formulated or selected to represent specific chemical types or performance levels, or both. They are usually supplied directly to a testing laboratory under code numbers to ensure that the laboratory is not influenced by prior knowledge of acceptable results in assessing the test results. The TMC determines which specific reference oil the laboratory shall test.

10.2.1.3  Unless specifically authorized by the TMC, do not analyze reference oils, either physically or chemically. Identification of reference oils by such analyses could undermine the confidentiality required to operate an effective reference oil system. Therefore, reference oils are supplied with the explicit understanding that they will not be subjected to analyses other than those specified in this procedure, unless specifically authorized by the TMC. If so authorized, prepare a written statement of the circumstances involved, the name of the person authorizing the analysis, and the data obtained; furnish copies of this statement to the TMC.

10.2.2  Reference Oil Test Frequency—Conduct reference oil tests according to the following frequency requirements:

10.2.2.1  For a given, calibrated test stand, conduct an acceptable reference oil test after no more than XX test starts have been conducted, or after TBD have elapsed, whichever occurs first.

10.2.2.2  After starting a laboratory reference oil test, non-reference oil tests may be started on any other calibrated test stand.

10.2.2.3  Reference oil test frequency may be adjusted due to the following reasons:

10.2.3  Procedural Deviations—On occasions when a laboratory becomes aware of a significant deviation from the test method, such as might arise during an in-house review or a TMC inspection, the laboratory and the TMC shall agree on an appropriate course of action to remedy the deviation. This action may include the shortening of existing reference oil calibration periods.

10.2.4  Parts and Fuel Shortages—Under special circumstances, such as industry-wide parts or fuel shortages, the Surveillance Panel may direct the TMC to extend the time intervals between reference oil tests. These extensions shall not exceed one regular calibration period.

10.2.5  Reference Oil Test Data Flow—To ensure continuous severity and precision monitoring, calibration tests are conducted periodically throughout the year. There may be occasions when laboratories conduct a large portion of calibration tests in a short period of time. This could result in an unacceptably large time frame when very few calibration tests are conducted. The TMC can shorten or extend calibration periods as needed to provide a consistent flow of reference oil test data. Adjustments to calibration periods are made such that laboratories incur no net loss (or gain) in calibration status.

10.2.6  Special Use of the Reference Oil Calibration System—The Surveillance Panel has the option to use the reference oil system to evaluate changes that have potential impact on test severity and precision. This option is only taken when a program of donated tests is not feasible. The Surveillance Panel and the TMC shall develop a detailed plan for the test program. This plan requires all reference oil tests in the program to be completed as close to the same time as possible, so that no laboratory/stand calibration is left in an excessively long pending status. In order to maintain the integrity of the reference oil monitoring system, each reference oil test is conducted so as to be interpretable for stand calibration. To facilitate the required test scheduling, the Surveillance Panel may direct the TMC to lengthen and shorten reference oil calibration periods within laboratories such that the laboratories incur no net loss (or gain) in calibration status.

10.2.7  Donated Reference Oil Test Programs—The Surveillance Panel is charged with maintaining effective reference oil test severity and precision monitoring. During times of new parts introductions, new or re-blended reference oil additions, and procedural revisions, it may be necessary to evaluate the possible effects on severity and precision levels. The Surveillance Panel may choose to conduct a program of donated reference oil tests in those laboratories participating in the monitoring system, in order to quantify the effect of a particular change on severity and precision. Typically, the Surveillance Panel requests its panel members to volunteer enough reference oil test results to create a robust data set. Broad laboratory participation is needed to provide a representative sampling of the industry. To ensure the quality of the data obtained, donated tests are conducted on calibrated test stands. The Surveillance Panel shall arrange an appropriate number of donated tests and ensure completion of the test program in a timely manner.

10.2.8  Reporting of Reference Oil Test Results—Report the results of all reference oil tests to the TMC according to the following directives:

10.2.8.1  Transmit results to the TMC within five days of test completion by way of electronic data transfer protocol as outlined in the Data Communication Committee, Electronic Test Report Transmission Model (ETRTM). The ETRTM can be obtained from the TMC.

10.2.8.2  If the test was conducted during a time extension permitted by the TMC, so indicate in the Comments section of the test report.

10.2.8.3  For an acceptable reference oil test, conducted following an unacceptable reference oil test, provide sufficient information in the Comments section of the test report to indicate how the problem was identified and corrected, insofar as possible, and how it was related to non-reference oil tests conducted during the period of time that the problem was being solved.

10.2.9  Evaluation of Reference Oil Test Results—The TMC evaluates the reference-oil test results for both operational validity and statistical acceptability. The TMC may consult with the test laboratory in case of difficulty, as follows:

10.2.9.1  Immediately upon receipt of the reference-oil test results from the test laboratory, the TMC evaluates the laboratories decision on operational validity. For operationally valid tests, the TMC then evaluates the pass/fail parameters according to the Sequence IVB Lubricant Test Monitoring System (TBD). If the test is judged acceptable, the reference oil code is disclosed by the TMC to the test laboratory. The TMC conveys to the test laboratory its preliminary findings based on the limited information available to them.

10.2.9.2  Subsequently, upon receipt of the information detailed in 13.1.1 the TMC reviews all reference-oil test results and reports to determine final test acceptability.

10.2.9.3  In the event the reference oil test is unacceptable, the test laboratory shall provide an explanation of the problem relating to the failure. If the problem is not obvious, all test-related equipment shall be re-checked. Following this re-check, the TMC assigns another reference oil for testing by the laboratory.

10.2.9.4  The TMC decides, with consultation as needed with industry experts (testing laboratories, members of the ASTM Technical Guidance Committee and of the Surveillance Panel, and so forth), whether the reason for any failure of a reference oil test is a false alarm, testing stand, testing laboratory, or industry-related problem. The Sequence IVB Surveillance Panel shall adjudicate all industry problems.

10.2.10  Status of Non-Reference Oil Tests Relative to Reference Oil Tests—Non-reference oil tests may proceed within a given laboratory during reference oil testing based upon the following:

10.2.10.1  During the time of conducting a reference oil test on one test stand, non-reference oil tests may be conducted on other previously calibrated stands. If the reference oil test is acceptable to the TMC, the non-reference oil tests shall be considered to have been run in a satisfactorily calibrated laboratory.

10.2.10.2  If a reference oil test is unacceptable, and it is determined that the problem is isolated to an individual test stand, consider other test stands to remain calibrated, and testing of non-reference oils may proceed on those other stands.

10.2.10.3  If a reference oil test is unacceptable, and it is determined that the problem is laboratory related, non-reference tests running during the problem period shall be considered invalid unless there is specific evidence to the contrary for each test.

10.2.11  Status of Test Stands Used for Non-Standard Tests—If a non-standard test is conducted on a previously calibrated test stand, conduct a reference oil test on that stand to demonstrate that it continues to be calibrated, prior to running standard tests.

10.3  Equipment Calibration:

10.3.1  Instrumentation Calibration—Perform a thorough recalibration adjustment of all instrumentation and transducers, including computer channels, according to the requirements that follow. Perform additional calibration checks whenever operational data indicates an abnormality. Standards used for instrumentation calibration shall be traceable to that country’s specific national standards organization. The accuracy of the standard shall be a minimum of four times better than the accuracy of the test stand instrumentation.

10.3.2  Dynamometer Torque Measurement—Scale the final readout of engine torque (N·m). Calibrate the force measurement and readout system with deadweights. Coolant flow through the dynamometer, reaction forces due to coolant plumbing, and brinnelled trunnion bearings of the dynamometer may affect calibration by temperature excursions of the dynamometer electronic force transducer. When calibrating, ensure the dynamometer coolant flow indicator is in the green (Refer to Figure 10 in Section E) and that the load cell temperature has been stabilized at 45°C ± 1°C for a minimum of one hour. The torque measurement accuracy shall be ± 0.2 N·m. Perform this calibration prior to every test start.

10.3.3 Instrument Calibration—Document all instrument calibrations.  Retain all calibration documentation for a minimum of 3 years.

10.3.3 Upon initial stand installation and every six months thereafter perform a full instrumentation calibration according to Table 9.

TABLE 9 Sequence IVB Instrument Calibrations to be Performed every 6 Months

Temperatures

Intake Air Temperature, °C

Engine Oil Gallery Temperature, °C

Engine Oil Sump Temperature, °C

Coolant Temperature Into Engine, °C

Coolant Temperature Out of Engine, °C

Fuel Rail Temperature, °C

Exhaust Gas Temperature, °C

Rocker Cover Coolant In Temperature, °C

Rocker Cover Coolant Out Temperature, °C

Test Cell Air Temperature, °C

Pressures

Crankcase Gas Pressure, kPa

Oil Gallery Pressure, kPa

Fuel Rail Pressure, kPa

Exhaust Pressure, kPa (absolute)

Intake Air Pressure, kPa

Intake Manifold Pressure, kPa (absolute)

Barometric Pressure, kPa (absolute)

Flows

Air Fuel Ratio, afr

Blowby Flow Rate, sl/min

Fuel Flow Rate, kg/h

General

Intake Air Humidity, grains/kg

Engine Speed, r/min

Engine Torque (N-m)

10.3.4  Humidity of Induction Air Calibration:

10.3.4.1  Calibrate the primary laboratory measurement system at each test stand every 6 months using a hygrometer with a minimum dew point accuracy of ± 0.55 °C at 16 °C. Locate the sample tap on the air supply line to the engine, between the main duct and 1000 mm upstream of the intake air cleaner. The calibration consists of a series of paired humidity measurements comparing the laboratory system with the calibration hygrometer. The comparison period lasts from 20 min to 2 h with measurements taken at intervals of (1 to 6) min, for a total of 20 paired measurements. The measurement interval shall be appropriate for the time constant of the humidity measurement instruments.

10.3.4.2  Verify that the flow rate is within the equipment manufacturer’s specification and that the sample lines are non-hygroscopic. Correct dew point hygrometer measurements to standard conditions (101.12 kPa) using the appropriate equation. Compute the difference between each pair of readings and calculate the mean and standard deviation of the twenty-paired readings. The absolute value of the mean difference shall not exceed 1.43 g/kg, and the standard deviation shall not be greater than 0.714 g/kg. If these conditions are not met, investigate the cause, make repairs, and recalibrate. Maintain calibration records for 2 years.

10.3.5  Profilometer Calibration—Follow the manufacturer’s instruction for calibration and verification checks of the profilometer. Calibrate the profilometer at least annually.

10.3.6  Keyence Measurement Device —Confirm the calibration of the Keyence measurement device at least once every 6 months.

11.  Procedure

11.1 Configure the engine for run-in and silicone pacification as follows:

11.1.1 First run engines have the correct engine break-in valve train components pre-installed.  Install the components listed in Annex XX and in the engine assembly manual sections XX.

11.1.2 If not a first run engine, then install stock intake camshaft, stock exhaust camshaft, and stock bucket lifters that have been designated for break-in use by the laboratory. Install the components listed in Annex XX and in the engine assembly manual sections ___XX__

Note 12: Pre-test measurements are not required on valve train components used for engine break-in purposes.

11.2 Configure the engine with stock valve train cover and active positive crankcase ventilation. Ensure the following connections are made.  Affix all ends of hoses with hose clamps.

    - Refer to Fig. 11.  Connect the stock PCV valve (A) to port (C) on the intake manifold (E) with the 19 mm diameter hose (B). This is the OEM supplied PCV hose.

    - Refer to Fig. 12.  Connect the port (A) on the valve train cover to port (B) on the intake air filter housing with the 15.8 mm  diameter Tygon hose (C).

    - Refer to Fig.13.  Connect the 8 mm steel-braided hose (B) to the quick-disconnect (A) on the modified oil fill cap.

Fig. 11: Routing of crankcase gases.  (A) stock PCV valve, (B) 19 mm diameter hose, (C) hose adapter to intake manifold, (D) stock valve train cover, (E) intake manifold. The 19 mm diameter hose is the OEM provided crankcase ventilation hose which is installed on new engines.

Fig. 12: Routing of fresh air.  (A) port to valve train cover, (B) port from intake air filter housing, (C) 15 mm diameter Tygon hose

Fig. 13: Crankcase pressure measurement point.  (A) oil fill cap modified with 5 mm diameter quick disconnect, (B) 8 mm diameter steel-braided hose to pressure transducer

  - Conduct the following external oil flush procedure.

11.3.1 Refer to Fig. 14.  Disconnect the supply and return lines from the remote oil filter housing adapter that is mounted on the engine.

Fig. 14: Engine-mounted oil filter housing adapter line connections.

(A) is return from oil cooler and (B) is supply to oil filter

    - Connect the supply and return lines to a portable oil cleaning flush cart of minimum 3.8 L capacity that is equipped with a circulation pump. Charge the flush cart with mineral spirits meeting the requirements of Specification D235, Type II, Class C for Aromatic Content (0 to 2 vol)%, Flash Point (61 °C, min) and Color (not darker than +25 on Saybolt Scale or 25 on Pt-Co Scale). Energize the flush cart pump and allow the mineral spirits to circulate for one h.

    - At the end of one h, de-energize the flush cart pump. Open both engine oil heat exchanger drain valves (see Fig. 15). Disconnect the supply and return lines from the flush cart.

Fig. 15: Engine Oil Heat Exchanger Drain Valve Locations

    - When the heat exchanger has completed draining, leave the heat exchanger drain valves open and connect both the supply and return lines to a clean, dry compressed air supply at 140 kPa. Allow compressed air to flow through the system for 15 min to dry the system.

    - Disconnect the supply and return lines from the compressed air source.

    - Connect the supply and return lines back to the engine-mounted oil filter housing adapter.

    - Close the heat exchanger drain valves, and remove the Oberg oil filter element for cleaning. Clear any debris retained in the Oberg oil filter element with mineral spirits and air dry. Re-install the Oberg oil filter element in the Oberg filter housing and secure the four retaining bolts as shown in Fig. 16.

Fig. 16: Oberg Filter Location

    - Dispose of the used mineral spirits following test laboratory practice.

  - Fill the engine with 3.0 L of break-in oil.  The break-in oil is ASTM TMC reference oil 1006-2.

  - Fill the engine coolant system and the rocker cover coolant system with a 70/30 mixture of de-ionized or distilled water and Havoline Extended Life DEX-COOL® concentrated anti-freeze.  Approximately 45 L of the 70/30 mixture is required to fill both systems.

  - With coolant flow control flow valve 50% open and the coolant system pressurized at 70 kPa, energize the engine coolant pump and allow the engine coolant to circulate for 5 min to remove air from the engine coolant system.

  - Using the coolant heater, gradually increase the coolant temperature to 50 °C over 15 min.

  - Soak the coolant temperature at 50 °C for 30 min.

  - With the dynamometer load control set point at 0 N·m and the throttle control set point at 10%, start the engine. As soon as the engine achieves 500 r/min increase the throttle control set point to 800 r/min while continuing to hold the dynamometer load control set point at 0 N·m.

  - When the engine achieves 800r/min begin the engine break-in schedule provided in Table 10.

Table 10 Engine Break-in Schedule

| Break-in Step no. | Duration min | Engine Speed r/min | Engine Load N·m | Gallery Oil Temperature °C | Coolant Out Temperature °C |
| --- | --- | --- | --- | --- | --- |
| 1 | 10 | 800 | 6.3 | 50 | 50 |
| 2 | 10 | 1600 | 6.3 | 55 | 50 |
| 3 | 10 | 2000 | 25.0 | 60 | 55 |
| 4 | 10 | 2400 | 25.0 | 65 | 60 |
| 5 | 10 | 2400 | 46.9 | 70 | 65 |
| 6 | 15 | 2800 | 46.9 | 75 | 70 |
| 7 | 15 | 3200 | 46.9 | 80 | 75 |
| 8 | 15 | 3200 | 46.9 | 85 | 80 |

11.10.1 The duration for each step includes the time to transition between set points.  Engine load should achieve specified value within 45 s, and engine speed should achieve specified value within 60 s.

    - The following parameters should be controlled to the specified set points for all break-in 	steps:

      - Fuel rail temperature  24 ± 3 ° C

      - Coolant temperature differential (Coolant out temp – Coolant In temp)  2 ± 2  ° C

      - Exhaust back pressure  103.5 ± 1.0 kPaA

      - Intake air pressure  0.07 ± 0.03 kPa

      - Intake air temperature  32 ± 2 ° C

      - Load cell temperature 45 ± 2 ° C

    - Following the completion of engine break-in step 11.8, establish the following oil sampling conditions as shown in Table 11.

Table 11 Oil Sampling Conditions

| Oil Sampling Conditions | Duration min | Engine Speed r/min | Engine Load N·m | Gallery Oil Temperature °C | Coolant Out Temperature °C |
| --- | --- | --- | --- | --- | --- |
|  | --- | 1000 | 10.0 | 50 | 50 |

    - When oil temperature has been reduced to 50°C, take a 240 mL purge sample immediately followed by a 30 mL oil sample.    Conduct ASTM D5185 Metals by ICP on the oil sample.  If the sample cannot be obtained within 10 min after achieving the 50°C oil temperature stop the engine as excessive idling time can be detrimental to engine run-in.

  - When the oil sample has been obtained in step 11.10.4, stop the engine. Disconnect the conditioned intake air supply. Return the purge sample to the engine via the factory oil fill cap. Allow the engine to rest for ten min.  Following the ten min rest period, measure and record the oil level by removing the dipstick and measuring from the bottom of the dipstick to the top of the oil film.  Record the length of the oil film in millimeters (see Fig. 17). Do not top off the oil level.

Fig. 17: Engine Oil Level Measurement

  - Start the engine, and continue to run the engine at the following conditions for engine aging as shown in Table 12:

Table 12 Engine Aging Conditions

| Aging Step no. | Duration (h) | Engine Speed r/min | Engine Load N·m | Gallery Oil Temp ° C | Coolant Out Temp ° C |
| --- | --- | --- | --- | --- | --- |
| Aging | 25 | 3000 | 50 | 90 | 88 |
| Sampling | --- | 1000 | 10 | 90 | 90 |

    - The duration for each step includes the time to transition between set points.  Engine load should achieve specified value in 45 s, and engine speed should achieve specified value in 60 s.

    - After every 25 h of aging, bring the engine down to the oil sampling conditions. If the sample cannot be obtained within 10 minutes after achieving the 50 °C oil temperature stop the engine as excessive idling time can be detrimental to engine run-in.

    - Take a 240 mL purge sample immediately followed by a 30 mL oil sample.  Return the purge sample to the engine via the factory oil fill cap.  Conduct ASTM D5185 Metals by ICP on the oil sample.

    - Stop the engine and allow the engine to rest for ten minutes.  Measure and record the oil level by removing the dipstick and measuring from the bottom of the dipstick to the top of the oil film.  Record the length of the oil film in millimeters. Do not top off the oil level.

    - Repeat steps 11.12 through 11.12.5 three additional times for a total aging time of one-hundred (100) hours.

    - Drain oil charge after the aging runs have completed.

11.13 Examine the results of the ASTM D5185 for high wear anomalies using Fe, Cu, and Al and to ensure the Si levels have plateaued. Also examine values of K as an indicator of coolant leaks. K values exceeding 15 ppm are suspicious and the engine should be evaluated for an internal coolant leak. If an internal coolant leak is confirmed, make repairs and repeat the run-in and aging procedure.

Note 18 – the example graph below (Fig. 18) is that of a new engine. Used engines with fresh cylinder heads will provide lower results.

Fig. 18: Example of Si Plateau (New Engine)

11.14 This completed the run-in and conditioning procedure.

12.  Calculation or Interpretation of Results

Note 19—The summary of results and calculations are recorded on the appropriate report form.

  - Preparing Pre-Test Lifters for Measurement:

    - Record the following information for each lifter:

      - Unique ID number engraved on the inside

      - Lifter grade

      - Position in the engine (including intake or exhaust side)

    - This information needs to be properly documented and updated as needed throughout the course of the test. An example of this documentation is shown in Table 13.

Table 13 - Example of Lifter Identification Documentation

| Intake Side |
| --- |
| Position | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
| ID No. |  |  |  |  |  |  |  |  |
| Grade |  |  |  |  |  |  |  |  |
| Exhaust Side |
| Position | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
| ID No. |  |  |  |  |  |  |  |  |
| Grade |  |  |  |  |  |  |  |  |

    - Clean the lifters using pentane or heptane and allow the solvent to evaporate before taking any measurements.

WARING-( Flammable, Health hazard.)  Available from petroleum solvent suppliers.

  - Measure the Pre-Test Lifter Weights:

    - Use a calibrated scale to measure and record the weight of the lifters.

    - These weights are to be reported to (4) decimal places in units of grams.

    - A Mettler AE200 Analytical Balance is recommended for these measurements.

  - Confirm Center Position of OHT Fixture:

    - Place a dimpled reference lifter on the OHT fixture (Fig. 19).

Fig. 19 - OHT Lifter Fixture

12.3.2 Preload the PDI stylus on the dimpled reference lifter so that its trace path is aligned with the dimple.

    - Perform a trace of the dimpled reference lifter through the dimple.

      - Move the stage or fixture as needed to find the lowest spot of the dimple.

      - This will effectively center the OHT fixture.

    - Save the trace of the dimpled lifter (through the lowest spot of the dimple) for future reference.

  - Pre-Test Lifter Profile Traces:

12.4.1 Profile traces are to be performed using a Precision Devices, Inc. (PDI) MicroAnalyzer 2000 profilometer that is using the Windows XP Version 3.6.15 operational software.

    - Confirm that the TOYOTA_BUCKET.NDT template file is being used.

    - Confirm the following settings on the Delimitation tab (Fig. 20):

      - The Find checkbox is selected and the filter width is set to 0.400mm under the (A) Left Edge field.

      - The Find checkbox is selected and the filter width is set to 0.400mm under the (B) Right Edge field.

      - The (C) Fix Truncation to Part Edges checkbox is selected.

      - The (D) ASME B46.1-2002 standard is being used.

Fig. 20 - Delimitation Tab Settings

    - Confirm the following settings on the Primary Extraction tab (Fig. 21):

      - The (E) Two-Point Line radial button is selected under the Form Removal field.

      - The filter cutoff is set to (F) 0.00µm under the Shortwave Cutoff field.

Fig.21 - Primary Extraction Tab Settings

    - Confirm the following settings on the Waviness Extraction tab (Fig. 22):

      - The (G) Gaussian w/o end removal radial button is selected under the Waviness Short Cutoff field.

      - The filter cutoff is set to (H) 0.800mm under the Waviness Short Cutoff field.

Fig. 22 - Waviness Extraction Tab Settings

    - Confirm the following settings on the Roughness Extraction tab (Fig. 23):

      - The (I) Gaussian w/o end removal radial button is selected under the Roughness Cutoff field.

      - The filter cutoff is set to (J) 0.800mm under the Roughness Cutoff field.

Fig. 23 - Roughness Extraction Tab Settings

    - Place the notched lifter on the OHT fixture with the wear surface positioned upward.

      - Profile traces are taken on two orthogonal lines as shown in Fig. 24

      - The x-axis intersects the center of the lifter face near the middle of the ID number engraved on the inside of the lifter.

      - The y-axis also passes through the center of the wear surface and is perpendicular to the x-axis.

Fig. 24 - Orthogonal Lines Used for Lifter Profile Measurements

    - Obtain a profile trace along one of the orthogonal lines by setting the OHT fixture in the “A” position.

Note 20: The current orthogonal axis of the OHT fixture can be easily identified by the position of the removable pin. For example, the OHT fixture is in the “A” position when the pin is near the “A” label on the black base plate.

    - Level the trace using the two-point line on the Primary Extraction tab.

      - Adjust the trace until its left and right edges are on the same horizontal level (as shown in Fig. 25).

Figure 25 - Leveling the Lifter Profile Trace

12.4.10 Save the pre-test lifter traces using the unique file name that is assigned by the PDI software.

      - Each laboratory is responsible for correlating this unique file name to the unique ID number assigned to each lifter.

      - The laboratory must also correlate this file name to the orthogonal axis used for the trace (“A” or “B” as indicated on the OHT fixture).

    - Record the pre-test Wt (waviness total) measurement (µm) with the waviness evaluation lines set on the extreme left-side and right-side edges of the waviness profile trace.

    - Rotate the lifter 90-degrees and obtain a profile trace along the “B” position of the OHT fixture.

      - Repeat Steps 12.4.9 through 12.4.11.

12.4.13 Acceptable and Unacceptable Lifter Traces:

    - .1 The acceptable pre-test waviness total for intake lifters is 2µm < Wt < 15µm.

      - The acceptable pre-test waviness total for exhaust lifters is 0µm < Wt.

      - Reject any lifter with surface irregularities that prevent it from having a continuous and symmetric crowned profile (examples shown in Fig. 26).

Fig. 26 – Examples of Unacceptable Lifter Traces

12.4.13.4 An example of a profile trace from an acceptable lifter can be found in Fig. 27.

Fig. 27 - Example of an Acceptable Lifter Trace

12.5   Preparing Post-Test Lifters for Measurement:

12.5.1 Clean the lifters using pentane or heptane and allow the solvent to evaporate before taking any measurements.

12.6 Measuring the Post-Test Lifter Weights:

    - Use a calibrated scale to measure and record the weight of the lifters.

      - These weights are to be reported to four decimal places in units of grams.

      - A Mettler AE200 Analytical Balance is recommended for these measurements.

12.7 Post-Test Lifter Profile Traces:

12.7.1Confirm that the TOYOTA_BUCKET.NDT template file is being used.

    - Confirm the following settings on the Delimitation tab (Fig. 20):

      - The Find checkbox is selected and the filter width is set to 0.400mm under the Left Edge field.

      - The Find checkbox is selected and the filter width is set to 0.400mm under the Right Edge field.

      - The Fix Truncation to Part Edges checkbox is selected.

      - The ASME B46.1-2002 standard is being used.

12.7.3 Confirm the following settings on the Primary Extraction tab (Fig. 21):

12.7.3.1 The Two-Point Line radial button is selected under the Form Removal field.

      - The filter cutoff is set to 0.00µm under the Shortwave Cutoff field.

    - Confirm the following settings on the Waviness Extraction tab (Fig. 22):

      - The Gaussian w/o end removal radial button is selected under the Waviness Short Cutoff field.

12.7.4.2           The filter cutoff is set to 0.800mm under the Waviness Short Cutoff field.

  - .5 Confirm the following settings on the Roughness Extraction tab (Fig. 23):

12.7.5.1The Gaussian w/o end removal radial button is selected under the Roughness Cutoff field.

12.7.5.2The filter cutoff is set to 0.800mm under the Roughness Cutoff field.

12.7.6 Obtain a profile trace along both of the lifter’s orthogonal axes (identified as “A” and “B” on the OHT fixture).

12.7.7 Save the post-test lifter traces using the unique file name that is assigned by the PDI 				software.

      - Each laboratory is responsible for correlating this unique file name to the unique ID number assigned to each lifter.

      - The laboratory must also correlate this file name to the orthogonal axis used for the trace (“A” or “B” as indicated on the OHT fixture).

    - Open the Comparator software.

12.7.8.1 All lifter profile comparisons are to be performed using Version 3.0.3 of the Comparator software.

12.7.9 Press the (K) Open Pre button on the main Comparator screen (Fig. 28).

      - Open the desired pre-test lifter trace document or documents.

12.7.10 Press the (L) Open Post button on the main Comparator screen as shown in Fig. 28.

    - Open the corresponding post-test lifter trace document or documents.

12.7.11 Use the (N) vertical adjustment lines and the (O) adjustment arrow buttons to level and align the pre-test and post-test traces (Fig. 28).

Fig. 28 - Main Comparator Screen

12.7.12 Press the (M) Compare button on the main Comparator screen as shown in Fig. 28.

12.7.12.1 Select (Q) Wear Comparison Waviness from the Comparison Type pull down menu of the document selection screen (Fig. 29).

12.7.12.2 Select the appropriate pre-test file (R) under the Pre Documents field.

12.7.12.3 Select the corresponding post-test file (S) under the Post Documents field.

12.7.12.4 Press the OK button.

Fig. 29 - Comparator Software Document Selection Screen

    - Record the required measurements (P, Fig. 28).

      - The Max Wear parameter is to be reported as z-diff.

  - Preparing Pre-Test Camshafts for Measurement:

    - Spray the camshafts with Stoddard solvent.

      - Use a non-metallic brush to thoroughly scrub the lobe surfaces.

      - Dry the camshaft with compressed air.

    - Visually inspect each camshaft lobe for defects or damage.

      - Examples of such defects or damage are shown in Fig. 30.

      - Reject any camshafts that have these abnormalities.

    - Record the unique identification number for both the intake and exhaust camshafts.

Fig. 30 - (A) Grinding Damage, (B) Pitting and (C) Shipping Damage

  - Pre-Test Camshaft Diameter Measurements:

12.9.1 Use a pin gage set to measure and record the diameter of all five oil feed holes on the journals of both the intake and exhaust camshafts (Fig. 31).

Fig. 31 - Measuring Oil Feed Hole Diameters Using a Pin Gage

    - Place the camshaft on a V-Block fixture.

      - The V-Blocks should be spaced so that each camshaft is supported on its 2nd and 5th journals.

    - Place a dial indicator next to the camshaft so that the spindle head is in contact with Journal #3 and is orthogonal to the axis of the camshaft (Fig. 32).

      - Rotate the camshaft until the dial indicator reads a minimum value.

      - Zero the dial indicator.

      - Rotate the camshaft again until the dial indicator reads a maximum value.

      - Record this dial indicator measurement as the run-out of Journal #3.

      - Repeat these steps to obtain the run-out measurement for Journal #4.

      - Average the run-out measurements of Journal #3 and Journal #4 to calculate the overall run-out of the camshaft.

12.9.3.7 Repeat these steps for the second camshaft.

Fig. 32 - Correct Positioning of Dial Indicator for Journal Run-Out Measurements

    - Reject any camshaft that has a run-out value that exceeds 0.04mm (0.00157-in).

    - Measure the diameter of Journal #1 with a 1-2 inch digital micrometer.

12.9.5.1 A Mitutoyo Model #293-722-10 digital micrometer is recommended for these measurements.

      - Measure the diameter of the journal along both its x-axis and y-axis (Fig. 33).

      - The x-axis passes through the center of the oil feed hole and intersects the center of the journal.

      - The y-axis is orthogonal to the x-axis.

      - Record the average of the x-axis and y-axis diameter measurements for Journal #1.

Fig. 33 - X and Y-Axis of Camshaft Journal

    - Use a 0-1 inch digital micrometer to measure the diameter of Journal #2 through Journal #5.

12.9.6.1 A Mitutoyo Model #293-721-10 digital micrometer is recommended for these measurements.

    - Reject any camshafts with journal diameters that exceed the following specifications (Table 14):

Table 14 - Camshaft Journal Diameter Specifications

| Item | Specified Condition |
| --- | --- |
| Journal #1 | 33.949 - 33.968 mm (1.3366 – 1.3372 in) |
| Journal #2 - #5 | – 22.965 mm (0.9035 – 0.9041 in) |

  - Pre-Test Camshaft Lobe Heel-to-Toe Height:

    - Place the camshaft on a V-Block fixture.

12.10.1.1 The V-Blocks should be spaced so that each camshaft is supported on its 2nd and 5th journals.

    - Calibrate a 1-2 inch snap gage using a cylindrical measurement standard.

      - A Mitutoyo Model #201-152 snap gauge with a Mitutoyo Model #ID-C112AEB digital indicator attachment is recommended for these measurements.

      - The cylindrical measurement standard used for intake camshaft lobe heel-to-toe measurements has a diameter of 36.725mm.

      - The cylindrical measurement standard used for exhaust camshaft lobe heel-to-toe measurements has a diameter of 39.500mm.

    - Set the base circle of the camshaft lobe on the anvil of the snap gage and engage the spindle (Fig. 34).

      - While keeping the snap gage centered on the lobe, slowly rotate the snap gage around the lobe until the digital indicator reads a maximum value.

      - Record this measurement as the heel-to-toe height for the specific camshaft lobe.

      - Repeat this measurement for the eight camshaft lobes on both the intake and exhaust camshafts.

Fig. 34 - Positioning the Snap Gage to Measure the Camshaft Heel-to-Toe Height

    - Reject any exhaust camshafts that have heel-to-toe heights that exceed the following specifications (Table 15):

      - There are currently no heel-to-toe height specifications for the intake camshafts.

Table 15 - Camshaft Heel-to-Toe Height Specifications

| Item | Specified Condition |
| --- | --- |
| Standard Lobe Height | 39.462-39.562 mm (1.554-1.558 in) |
| Minimum Lobe Height | (1.550 in) |

12.11 Pre-Test Camshaft Lobe Surface Profile Traces:

12.11.1 Profile traces are to be performed using a Precision Devices, Inc. (PDI) MicroAnalyzer 2000 profilometer that is using the Windows XP Version 3.6.15 operational software.

    - Confirm that the TOYOTA_CAM.NDT template file is being used.

    - Confirm the following settings on the Delimitation tab:

      - The Find checkbox is selected and the filter width is set to 0.400mm under the Left Edge field.

      - The Find checkbox is selected and the filter width is set to 0.400mm under the Right Edge field.

      - The Fix Truncation to Part Edges checkbox is selected.

      - The ASME B46.1-2002 standard is being used.

    - Confirm the following settings on the Primary Extraction tab:

      - The Two-Point Line radial button is selected under the Form Removal field.

      - The filter cutoff is set to 0.00µm under the Shortwave Cutoff field.

    - Confirm the following settings on the Waviness Extraction tab:

      - The Gaussian radial button is selected under the Waviness Short Cutoff field.

      - The filter cutoff is set to 0.800mm under the Waviness Short Cutoff field.

    - Confirm the following settings on the Roughness Extraction tab:

      - The Gaussian radial button is selected under the Roughness Cutoff field.

      - The filter cutoff is set to 0.800mm under the Roughness Cutoff field.

12.11.7 Place the camshaft on a V-Block fixture.

12 11.7 1 The V-Blocks should be spaced so that each camshaft is supported on its 2nd and 5th journals.

    - Move the stylus until it is over the camshaft lobe that is to be measured.

      - Rotate the camshaft until the nose of the lobe is pointed toward the stylus.

      - Lower the stylus onto the surface of the camshaft lobe.

12.11.8.3 Slowly rotate the camshaft until the stylus reaches the point of highest displacement.

    - Obtain a trace that spans the width of the camshaft lobe.

      - Report the roughness average (Ra), skew (Rsk) and waviness (Wt).

      - Save the pre-test traces using the appropriate file naming convention

      - Repeat the measurements for all eight camshaft lobes for both the intake and exhaust camshafts.

# imaging pre-test intake and exhaust lifters (AUTO STITCHING)

12.12.1 Establish the Correct Viewer Software Settings:

      - Open the Keyence VR-3000 Viewer software.

      - Select Options from the Tools pull-down menu at the top of the screen.

12.12.1.2. (a)This will cause the Options pop-up menu to appear.

12.12.1.3 Under the Options pop-up menu (Fig. 35):

12.12.1.3 (a)Select the (A) 3D measurement – Auto stitching option from the left-side of the pop-up menu.

12.12.1.3.(b)Uncheck the (B) Perform auto stitching after measurement and then 		   automatically open stitched image checkbox if needed.

12.12.1.3.(c)Check the (C) Auto-adjust position when stitching checkbox.

12.12.1.3.(d)Click the (D) OK button when complete.

Fig. 35 – Options [3D measurement – Auto stitching] Pop-Up Menu

      - Under the Change Magnification section located in the upper toolbar (Fig. 36):

12.12.1.4.(a)Select Low Mag Cam from the left pull-down menu.

12.12.1.4.(b)Select 25x from the Magnification pull-down menu.

Fig. 36 - Change Magnification Toolbar

12.12.1.5. Under the Expert Mode option of the 3D Measure tab (Fig. 37):

12.12.1.5.(a)Select the (A) Meas settings button.

12.12.1.5.(b)Select (B) Standard from the Mode pull-down menu.

12.12.1.5.(c)Select (C) Both Sides from the Measurement direction pull-down menu.

12.12.1.5.(d)Select the (D) Auto radial button under the Adjust brightness for measurement section.

12.12.1.5.(e)Select (E) Auto from the Stitching pull-down menu. This should be set to Manual for Keyence units with a manual stage.

12.12.1.5.(f)Check the (F) Enable AF checkbox.

Fig. 37 - Measurement Settings under 3D Measure Tab

12.12.1.6. Under the Expert Mode option of the 3D Measure tab (Fig. 38):

        - Select the (A) Stitching settings button.

        - Select the (B) Simple radial button under the Area specification method section.

Fig. 38 - Stitching Settings under 3D Measurement Tab

12.12.2 Take the Lifter Measurements:

12.12.2.1 All pre-test lifter images must be captured and post-processed before the test can be started.

12.12.2.2 Wipe the top of the lifter until it is free of residue, smudges and dust.

12.12.2.3 Place the lifter on the scanning table with the wear surface facing upwards and the serial number (inscribed underneath) facing the front of the macroscope.

12.12.2.4Adjust the lifter’s position on the scanning table until it is in view on the screen.

12.12.2.5 Select an area on the image of the lifter and double-click the left mouse button to perform an auto focus.

12.12.2.6 Position the camera near the center of the lifter and then pan the camera to the left until there is approximately1.0±0.25 in. of negative space between the left lifter edge and left-hand side of the screen (Fig. 39). Click and hold the right mouse button to drag the lifter image as needed, or use the arrows on the XY stage menu to move the camera.

Fig. 39– Negative Space between Left Lifter Edge and Left-Side of Screen

12.12.2.6.1. From this position, pan the camera upward until there is approximately 1.25±0.25 in. of negative space between the top lifter edge and the top of the screen. The distances specified above are based on a 1920X1080 screen resolution. The ideal amount of negative space may vary based on the monitor settings and screen resolution.

12.12.2.6.2. Press the (A) Add button to capture the image within the view of the Stitching area (Fig.40).

12.12.2.6.3. Repeat these steps to establish an image of the bottom right-hand side of the lifter.

12.12.2.6.4. Capturing images of the upper left-hand and bottom right-hand corners of the lifter (as shown in Fig. 40), or images of the lifter along its x-axis and y-axis edges, are sufficient to establish the stitching area.

Fig. 40 - Establishing the Stitching Area

12.12.2.7 Reference Addendum A for Keyence units that are equipped with a manual stitching table.

12.12.2.8 Click the Measure button in the lower right-hand side of the screen once the boundaries of the lifter are established.

12.12.2.9 Click the (A) Auto position adjustment button that appears on the right of the Stitching screen after the image capture is complete (Fig. 41).

12.12.2.10 Click the (B) Execute Stitching button that also appears on the right of the Stitching screen (Fig. 41).

Fig. 41 - Right-Side of Stitching Screen after Image Capture

12.12.2.11Under the Save Stitched Results prompt, select 1200X1200 from the Image size to save pull-down menu (Fig. 42).

12.12.2.11.(a)The user may be required to select the option from the pull-down menu that is the closest to the 1200X1200 size.

Fig. 42- Save Stitched Results Field

12.12.2.12 Click the Save stitched result button under the Save Stitched Results field (Fig. 42).

12.12.2.12.(a) Enter the desired file name in the Save As pop-up menu and press the Save button.

12.12.2.12.(b) Use a different naming convention for the pre-test and post-test image files.

12.12.2.13Select the Use template file checkbox on the Open the stitched result, which was saved, in Analyzer prompt.

12.12.2.13.(a) Click the Yes button.

12.12.2.14 Select the IVB Lifter Volume Temp clr Exhaust_rev2.ztp template file from the Open prompt.

12.12.2.14.(a) Press the Open button on the Open prompt.

12.12.2.15 Collecting Images

12.12.2.15.(a)The user also has the option of collecting all of the images first, and then post-processing them at the same time using the Analyzer software.

12.12.2.15.(b)This can be done in the Analyzer software by selecting the Open with Template option under the File pull-down menu.

12.12.2.15.(c)Select the IVB Lifter Volume Temp_clr Exhaust_rev2.ztp template file from the Open prompt.

12.12.2.15.(d)Use the Shift button to select all of the lifter images that are to have the template applied.

12.12.2.15.(e) Click the Ok button.

12.12.3 Post-Process the Images Using the Analyzer Software:

12.12.3.1Click the Ref plane button located in the upper toolbar of the main Analyzer Software screen (Fig. 43).

Fig. 43 - Reference Plane Button

12.12.3.2 Click the (A) Area settings button near the top of the Reference plane settings pop-up menu (Fig. 44).

Fig. 44 - Reference Plane Settings Pop-Up Menu

12.12.3.3 Click the (A) Import area button near the lower right-side corner of the Area settings pop-up menu (Fig. 45).

Fig. 45 - Area Settings Pop-Up Menu

      - Select the IVB leveling ring Exhaust_rev2.zri file from the Open prompt.

12.12.3.4.(a)Press the Open button.

12.12.3.5 It is very important to confirm that the leveling ring is concentric with the lifter face.

12.12.3.6 Click the (B) Ok button (Figure 45) on the Area settings pop-up menu.

12.12.3.7 Click the (B) Confirm button (Figure 44) the Reference plane settings pop-up menu.

12.12.3.8 Click the Ok button on the Reference plane settings (confirmation) pop-up menu.

      - Select the Volume & area 1 tab near the top of the Analyzer Software screen.

12.12.3.9. (a)The data in the volume vs. area ratio table (Fig. 46) needs to be reported in the final test report as the pre-test volume.

Fig. 46 - Critical Data on Volume & Area 1 Tab

12.12.3.10 Select the Save As option from the File pull-down menu near the top of the screen.

12.12.3.10.(a) Save this file using a different name than the original image in the event that the template application process is unsuccessful.

12.12.3.11 Click the To Viewer button.

      - Click the Exit Stitching button on the Viewer Software screen.

12.12.3.12.(a)Click the Yes button at the Ok to proceed prompt.

      - Click the Reset button underneath the Stitching area.

12.12.3.13.(a) Click the Yes button at the Delete current stitch area settings prompt.

# 12.13 imaging POST-test intake and exhaust lifters (AUTO STITCHING)

    - Establish the Correct Viewer Software Settings:

12.13.1.1 Repeat Steps 12.12.1.1 through 12.12.1.4.

      - Under the Expert Mode option of the 3D Measure tab (reference Fig. 47):

      - (a) Select the (A) Meas settings button.

      - (b) Select Glare Removal from the Mode pull-down menu (reference Item B).

      - (c)Select (C) Both Sides from the Measurement direction pull-down menu.

12.13.1.5.(d)Select the Manual radial button under the Adjust brightness for measurement section (reference Item D).

12.13.1.6(e) Select (E) Auto from the Stitching pull-down menu. This should be set to Manual for Keyence units with a manual stage.

12.13.1.7.(f) Select the (F) Enable AF checkbox.

12.13.3 Set the (A) Left side projection and (B) Right side projection sliders to 300.0ms (Fig. 47).

Fig. 47 - Projection Sliders to Adjust Brightness

    - Under the Expert Mode option of the 3D Measure tab:

12.13.4.1 Select the Stitching settings button.

12.13.4.2 Select the Simple radial button under the Area specification method section.

  - Take the Lifter Measurements:

12.14.1 Repeat Steps 12.12.2.1 through 12.12.2.14.

12.15 Post-Process the Images Using the Analyzer Software:

12.15.1 Repeat Steps 12.12.3.1 through 12.12.3.9.

    - Select the Volume & area 2 tab near the top of the Analyzer Software screen.

12.15.2.1 The data in the volume vs. area ratio table (Fig. 48) needs to be reported in the final test report as the post-test high spot volume.

Fig. 48 - Critical Data on Volume & Area 2 Tab

12.15.3 Select the Save As option from the File pull-down menu near the top of the screen.

12.15.3.1 Save this file using a different name than the original image in the event that the template application process is unsuccessful.

12.15.4 Click the To Viewer button.

    - Click the Exit Stitching button on the Viewer Software screen.

12.15.5.1 Click the Yes button at the Ok to proceed prompt.

    - Click the Reset button underneath the Stitching area.

12.15.6.1 Click the Yes button at the Delete current stitch area settings prompt.

# calculating lifter wear

12.16.1 Use the Following Equation to Calculate End of Test Wear for Each Lifter:

(2)

# instructions for keyence units with manual stage

    - Special Instructions for Manual Stages:

12.17.1.1 Position the camera near the center of the lifter and then pan the camera 	to the left until there is approximately 1.0±0.25 in. of negative space between the 	left lifter edge and left-hand side of the screen. Click and hold the right mouse 	button to drag the lifter image as needed, or use the arrows on the XY stage menu 	to move the camera.

12.17.1.1.(a) From this position, pan the camera upward until there is 	approximately 1.25±0.25 in. of negative space between the top lifter edge 	and the top of the screen.

12.17.1.1.(b) Press the Add button to capture the image within the view of 	the Stitching area.

12.17.1.2 When prompted to move to the next stitching position, pan the camera to 	the right of the lifter.

12.17.1.2.(a) Carefully align the transparent image with the opaque image.

12.17.1.3 Repeat Steps 12.17.1 and 12.17.2 until (12) images are captured.

12.17.1.4 The complete image of the lifter face should be (3) images wide by (4) 			images tall.

# 12.18 CALIBRATION INSTRUCTIONS

    - Recommended Calibration and Verification Frequencies:

12.18. 1.1 The manufacturer recommends verifying the calibration every time 			the macroscope is used.

12.18.1.2 The manufacturer recommends updating the calibration approximately 		one time per month.

12.18.3 All of the manufacturer’s calibration documentation can be found in the 		Viewer Software Reference Manual VR-H1V.

12.18.4 The macroscope must be allowed to warm-up for 1-h before performing a calibration or verification check.

    - XY-Calibration of Macroscope:

12.18.2.1 Place the calibration block on the stage plate. The calibration block is 			available through Keyence (P/N OP-87710).

12.18.2.2Adjust the camera position until the right-side of the calibration block 	is in view on the screen.

12.18.2.3 Select Calibration ( XY calibration from the Maintenance option in 	the Tools pull-down menu.

12.18.2.3.(a) Follow the prompts on the screen to complete the calibration.

    - Z-Calibration of Macroscope:

12.18.3.1 Adjust the camera position until the left-side of the calibration block is 	in view on the screen.

12.18.3.2 Select Calibration ( Z calibration from the Maintenance option in the Tools pull-down menu.

12.18.3.2.(a) Follow the prompts on the screen to complete the 			   calibration.

    - Verifying the Calibration:

12.18.4.1 Select Verification ( Measure width 20 times from the 			               Maintenance option in the Tools pull-down menu.

12.18.4.1.(a)Follow the prompts on the screen to complete the calibration.

12.18.4.2 Select Verification ( Measure height 20 times from the 	Maintenance option in the Tools pull-down menu.

12.18.4.2.(a)Follow the prompts on the screen to complete the calibration.

13.  Test Report

13.1 For reference oil results, use the standardized report form set available from the ASTM TMC and data dictionary for reporting test results and for summarizing operational data.

NOTE 21—Report the non-reference oil test results on these same forms if the results are intended to be submitted as candidate oil results against a specification.

13.1.1 Fill out the report forms according to the formats shown in the data dictionary.

13.1.2 Transmit results to the TMC within 5 working days of test completion.

13.1.3 Transmit the results electronically as described in the ASTM Data Communications Committee Test Report Transmission Model (Section 2 — Flat File Transmission Format) available from the ASTM TMC. Upload files via the TMC’s website.

13.2 Report all reference oil test results, whether aborted, invalidated, or successfully completed, to the TMC. Comment [TB1]: Insert appropriate Note number.

13.3 Deviations from Test Operational Limits—Report all deviations from specified test operational limits.

13.4 Precision of Reported Units—Use the Practice E29 rounding-off method for critical pass/fail test result data. Report the data to the same precision as indicated in data dictionary.

13.5 In the space provided, note the time, date, test hour, and duration of any shutdown or off-test condition. Document the outcome of all prior reference oil tests from the current calibration sequence that were operationally or statistically invalid.

13.6 If a calibration period is extended beyond the normal calibration period length, make a note in the comment section and attach a written confirmation of the granted extension from the TMC to the test report. List the outcomes of previous runs that may need to be considered as part of the extension in the comment section.

14.  Precision and Bias

14.1  Precision:

14.1.1  Test precision is established on the basis of reference oil test results (for operationally-valid tests) monitored by the TMC. The data are reviewed semiannually by the Sequence IVA Surveillance Panel. Contact TMC for the current industry data.

14.1.1.1  Intermediate Precision Conditions—Conditions where test results are obtained in the same laboratory with the same test method using the same test oil, with changing conditions such as operators, measuring equipment, test stands, and time between tests.

Note 21—“Intermediate precision” is the appropriate term for this test method, rather than “repeatability,” which defines more rigorous within-laboratory conditions.

14.1.1.2  Intermediate Precision Limit (i.p.)—The difference between two results obtained under intermediate precision conditions that would in the long run, in the normal and correct conduct of the test method, exceed the value show in Table 16, in only one case in twenty. When only a single test result is available, the Intermediate Precision Limit can be used to calculate a range (test result ± Intermediate Precision Limit) outside of which a second test result would be expected to fall about one time in twenty.

TABLE 16 Test PrecisionA

| Test Result | Intermediate PrecisionB | ReproducibilityC |
| --- | --- | --- |
|  | si.p.D | i.p.E | sRD | RE |
| Average Wear (μm) | 12.52 | 35.06 | 12.56 | 35.17 |

A Based on results obtained on ASTM reference oil 1006-2 from Feb. 20, 2002 to May 18, 2005.

B See 14.1.1.2.

C See 14.1.1.4.

D s = standard deviation.

E On the basis of test error alone, the difference, in absolute value, between two test results will be expected to exceed this value only about 5 % of the time.

14.1.1.3  Reproducibility Conditions—Conditions where test results are obtained with the same test method using the same test oil in different laboratories with different operators using different equipment.

14.1.1.4  Reproducibility Limit (R)—The difference obtained under reproducibility conditions that would in the long run, in the normal and correct conduct of the test method, exceed the value shown in Table 16, in only one case in twenty. When only a single test result is available, the Reproducibility Limit can be used to calculate a range (test result ± Reproducibility Limit) outside of which a second test result would be expected to fall about one time in twenty.

14.2  Bias—Bias is determined by applying an acceptable statistical technique to reference oil test results and when a significant bias is determined, a severity adjustment is permitted for non-reference oil test results. This technique is described in TMC 94-200.10

15.  Keywords

15.1 Cam lobe wear; crankcase oils; lubricants; valve-train wear

ANNEXES

(Mandatory Information)

A1.  ASTM TEST MONITORING CENTER: ROLE

A1.1  Nature and Functions of the ASTM Test Monitoring Center (TMC)—The TMC is a non-profit organization located in Pittsburgh, Pennsylvania and is staffed to: administer engineering studies; conduct laboratory inspections; perform statistical analyses of reference oil test data; blend, store, and ship reference oils; and provide the associated administrative functions to maintain the referencing calibration program for various lubricant tests as directed by ASTM Subcommittee D02.B0 and the ASTM Executive Committee. The TMC coordinates its activities with the test sponsors, the test developers, the surveillance panels, and the testing laboratories. Contact TMC through the TMC Director at:

ASTM Test Monitoring Center

6555 Penn Avenue

Pittsburgh, PA 15206-4489

www.astmtmc.cmu.edu

A1.2  Rules of Operation of the ASTM TMC—The TMC operates in accordance with the ASTM Charter, the ASTM Bylaws, the Regulations Governing ASTM Technical Committees, the Bylaws Governing ASTM Committee D02, and the Rules and Regulations Governing the ASTM Test Monitoring System.

A1.3  Management of the ASTM TMC—The management of the Test Monitoring System is vested in the Executive Committee elected by Subcommittee D02.B0. The Executive Committee selects the TMC Director who is responsible for directing the activities of the TMC.

A1.4  Operating Income of the ASTM TMC—The TMC operating income is obtained from fees levied on the reference oils supplied and on the calibration tests conducted. Fee schedules are established by the Executive Committee and reviewed by Subcommittee D02.B0.

# A2.  ASTM TEST MONITORING CENTER: CALIBRATION PROCEDURES

A2.1  Reference Oils—These oils are formulated or selected to represent specific chemical, or performance levels, or both. They are usually supplied directly to a testing laboratory under code numbers to ensure that the laboratory is not influenced by prior knowledge of acceptable results in assessing test results. The TMC determines the specific reference oil the laboratory shall test.

A2.1.1  Reference Oil Data Reporting – Test laboratories that receive reference oils for stand calibration shall submit data to the TMC on every sample of reference oil they receive.  If a shipment contains any missing or damaged samples, the laboratory shall notify the TMC immediately.

A2.2  Calibration Testing:

A2.2.1  Full-scale calibration testing shall be conducted at regular intervals. These full-scale tests are conducted using coded reference oils supplied by the TMC. It is a laboratory's responsibility to keep the on-site reference oil inventory at or above the minimum level specified by the TMC test engineers.

A2.2.2  Test Stands Used for Non-Standard Tests—If a non-standard test is conducted on a previously calibrated test stand, the laboratory shall conduct a reference oil test on that stand to demonstrate that it continues to be calibrated, prior to running standard tests.

A2.3  Reference Oil Storage—Store reference oils under cover in locations where the ambient temperature is between -10 °C and +50 °C.

A2.4  Analysis of Reference Oil—Unless specifically authorized by the TMC, do not analyze TMC reference oils, either physically or chemically. Do not resell ASTM reference oils or supply them to other laboratories without the approval of the TMC. The reference oils are supplied only for the intended purpose of obtaining calibration under the ASTM Test Monitoring System. Any unauthorized use is strictly forbidden. The testing laboratory tacitly agrees to use the TMC reference oils exclusively in accordance with the TMC’s published Policies for Use and Analysis of ASTM Reference Oils, and to run and report the reference oil test results according to TMC guidelines. Additional policies for the use and analysis of ASTM Reference Oils are available from the TMC.

A2.5  Conducting a Reference Oil Test—When laboratory personnel are ready to run a reference calibration test, they shall request an oil code via the TMC website.

A2.6  Reporting Reference Oil Test Results—Upon completion of the reference oil test, the test laboratory transmits the data electronically to the TMC, as described in Section 13. The TMC reviews the data and contacts the laboratory engineer to report the laboratory's calibration status. All reference oil test results, whether aborted, invalidated, or successfully completed, shall be reported to the TMC.

A2.6.1 All deviations from the specified test method shall be reported.

# A3. ASTM TEST MONITORING CENTER: MAINTENANCE ACTIVITIES

A3.1  Special Reference Oil Tests—To ensure continuous severity and precision monitoring, calibration tests are conducted periodically throughout the year. Occasionally, the majority or even all of the industry’s test stands will conduct calibration tests at roughly the same time. This could result in an unacceptably large time frame when very few calibration tests are conducted. The TMC can shorten or extend calibration periods as needed to provide a consistent flow of reference oil test data. Adjustments to calibration periods are made such that laboratories incur no net loss or gain in calibration status.

A3.2  Special Use of the Reference Oil Calibration System—The surveillance panel has the option to use the reference oil system to evaluate changes that have potential impact on test severity and precision. This option is only taken when a program of donated tests is not feasible. The surveillance panel and the TMC shall develop a detailed plan for the test program. This plan requires all reference oil tests in the program to be completed as close to the same time as possible, so that no laboratory/stand calibration status is left pending for an excessive length of time. In order to maintain the integrity of the reference oil monitoring system, each reference oil test is conducted so as to be interpretable for stand calibration. To facilitate the required test scheduling, the surveillance panel may direct the TMC to lengthen and shorten reference oil calibration periods within laboratories such that the laboratories incur no net loss or gain in calibration status. To ensure accurate stand, or laboratory, or both severity assessments, conduct non-reference oil tests the same as reference oil tests.

A3.3  Donated Reference Oil Test Programs—The surveillance panel is charged with maintaining effective reference oil test severity and precision monitoring. During times of new parts introductions, new or re-blended reference oil additions, and procedural revisions, it may be necessary to evaluate the possible effects on severity and precision levels. The surveillance panel may choose to conduct a program of donated reference oil tests in those laboratories participating in the monitoring system, in order to quantify the effect of a particular change on severity and precision. Typically, the surveillance panel requests its panel members to volunteer enough reference oil test results to create a robust data set. Broad laboratory participation is needed to provide a representative sampling of the industry. To ensure the quality of the data obtained, donated tests are conducted on calibrated test stands. The surveillance panel shall arrange an appropriate number of donated tests and ensure completion of the test program in a timely manner.

A3.4  Intervals Between Reference Oil Tests—Under special circumstances, such as extended downtime caused by industry-wide parts or fuel shortages, the TMC may extend the intervals between reference oil tests.

A3.5  Introducing New Reference Oils—Reference oils produce various results. When new reference oils are selected, participating laboratories will be requested to conduct their share of tests to enable the TMC to recommend industry test targets. ASTM surveillance panels require a minimum number of tests to establish the industry test targets for new reference oils.

A3.6  TMC Information Letters—Occasionally it is necessary to revise the test method, and notify the test laboratories of the change, prior to consideration of the revision by Subcommittee D02.B0. In such a case, the TMC issues an Information Letter. Information Letters are balloted semi-annually by Subcommittee D02.B0, and subsequently by D02. By this means, the Society due process procedures are applied to these Information Letters.

A3.6.1   Issuing Authority—The authority to issue an Information Letter differs according to its nature. In the case of an Information Letter concerning a part number change which does not affect test results, the TMC is authorized to issue such a letter. Long-term studies by the surveillance panel to improve the test procedure through improved operation and hardware control may result in the issuance of an Information Letter. If obvious procedural items affecting test results need immediate attention, the test sponsor and the TMC issue an Information Letter and present the background and data supporting that action to the surveillance panel for approval prior to the semiannual Subcommittee D02.B0 meeting.

A3.7  TMC Memoranda—In addition to the Information Letters, supplementary memoranda are issued. These are developed by the TMC and distributed to the appropriate surveillance panel and participating laboratories. They convey such information as batch approvals for test parts or materials, clarification of the test procedure, notes and suggestions of the collection and analysis of special data that the TMC may request, or for any other pertinent matters having no direct effect on the test performance, results, or precision and bias.

# A4. ASTM TEST MONITORING CENTER: RELATED INFORMATION

A4.1  New Laboratories—Laboratories wishing to become part of the ASTM Test Monitoring System will be requested to conduct reference oil tests to ensure that the laboratory is using the proper testing techniques. Information concerning fees, laboratory inspection, reagents, testing practices, appropriate committee membership, and rater training can be obtained by contacting the TMC Director.

A4.2  Information Letters: COTCO Approval—Authority for the issuance of Information Letters was given by the committee on Technical Committee Operations in 1984, as follows: “COTCO recognizes that D02 has a unique and complex situation. The use of Information Letters is approved providing each letter contains a disclaimer to the affect that such has not obtained ASTM consensus. These Information Letters should be moved to such consensus as rapidly as possible.”

A4.3  Precision Data—The TMC determines the precision of test methods by analyzing results of calibration tests conducted on reference oils. Precision data are updated regularly. Current precision data can be obtained from the TMC

AX.  PARTS LIST

A2.1  This annex illustrates the parts needed for the Sequence IVA test (Table A2.1).

TABLE A2.1 Parts List

| Section | Description | Part Number | Contents | Supplier |
| --- | --- | --- | --- | --- |
| 6.1 | Bare Engine Assembly | A0102-76P01 | Engine Block / Head / Valvetrain AssemblyA | Nissan North America, Inc. |
| 6.4.1.3 | Test Kit | 13000-40F85 | Camshaft Assembly (1) Rocker Shaft (2) Rocker Arms (12) Oil Filter Assembly (3) Spark Plug (4) | Nissan North America, Inc. |
| 6.4.1.4 | Head AssemblyB | A1040-40F80 | with Valves and Springs without Camshaft, Rocker Arms | Nissan North America, Inc. |
| 6.2.7 | Oil Cooler | 21305-03E00 | Engine Oil Cooler | Nissan North America, Inc. or |
|  | Oil Cooler | OHTKA24-006-1 | Engine Oil Cooler | OH Technologies, Inc., and |
|  | Adapter | OHTKA24-005-1 | Adapter for OH oil cooler | OH Technologies, Inc. |
|  | Mounting Stud | OHTKA24-007-1 | Mounting Stud for OH Oil Cooler | OH Technologies, Inc. |
| 6.4.1.4 | Engine Valve Regrind Kit | A1042-10C2E | Head Gasket and Seals | Nissan North America, Inc. |
| 6.4.1.2 | Test Stand Kit No. 1 | A0001-76P25 |  | Nissan North America, Inc. |
| 6.4.1.2 | Test Stand Kit No. 2 | A1001-40F25 |  | Nissan North America, Inc. |
| 6.4.1.2 | Test Stand Kit No. 3 | B4010-40F26 |  | Nissan North America, Inc. |
| 6.4.1.2 | Test Stand Kit No. 4 | 14004-F4003 |  | Nissan North America, Inc. |
| 6.3.9 | Jacketed Rocker Cover | TEI-NIVAWCR-020 | Aluminum Jacketed Rocker Cover | Test Engineering, Inc. |
| 6.2.9 | Modified Wiring Harness | OHTKA24-002-1 | Modified Harness for ECM | OH Technologies, Inc. |
| 6.3.4.2 | Air Filter Assembly | 16500-86G50KT | Air Filter Housing and Element | Nissan North America, Inc. |
| 6.5.1 | Cam Angle Encoder | NIVACWM010 |  | Test Engineering, Inc. |
|  | Cylinder Head Calibration Apparatus |  |  | OH Technologies, Inc. |
| 7.4.2 | Silicone Gasket Maker | 999MP-A7007 | RTV sealant | Nissan dealer |
| 7.2 | Test Fuel | KA24E | KA24E (dyed green) | Dow Chemical |
| 7.3.1 | Break-In Oil | TMC 926-2 | TMC926-2 | ASTM Test Monitoring Center |

A Can be used for 48 tests; cylinder head included with assembly can be used for 24 tests.

B Can be used for 24 tests.

AX.  FUELS SPECIFICATION INFORMATION

A4.1This annex provides information on the test fuel and engine coolant used in the Sequence IVA test procedure.

A4.1.1 KA24E Test Fuel (Fig.A4.1).

| Sales specification |
| --- |
| Telephone: (800) 969-2542 | FAX: (281) 457-1469 |
| Johann Haltermann Ltd. |
| PRODUCT: | KA24E TEST FUEL |
| PRODUCT CODE: | HF0008 |
| TEST | METHOD | UNITS | SPECIFICATIONS |
| MIN | TARGET | MAX |
| Distillation - IBP | ASTM D86 | °F | 75 | 95 |
| 5% | °F |
| 10% | °F | 120 | 135 |
| 20% | °F |
| 30% | °F |
| 40% | °F |
| 50% | °F | 200 | 230 |
| 60% | °F |
| 70% | °F |
| 80% | °F |
| 90% | °F | 300 | 325 |
| 95% | °F |
| Distillation - EP | °F | 385 | 415 |
| Recovery | vol % | Report |
| Residue | vol % | Report |
| Loss | vol % | Report |
| Gravity | ASTM D4052 | °API | 58.7 | 61.2 |
| Density | ASTM D4052 | kg/l | 0.734 | 0.744 |
| Reid Vapor Pressure | ASTM D5191 | psi | 8.8 | 9.2 |
| Carbon | ASTM D5291 | wt fraction | 0.8580 | 0.8667 |
| Carbon | ASTM D3343 | wt fraction | Report |
| Sulfur | ASTM D2622 | wt % | 0.01 | 0.04 |
| Lead | ASTM D3237 | g/gal | 0.05 |
| Oxygen | ASTM D4815 | wt % | 0.2 |
| Composition, aromatics | ASTM D1319 | vol % | 35.0 |
| Composition, olefins | ASTM D1319 | vol % | 5.0 | 10.0 |
| Composition, saturates | ASTM D1319 | vol % | Report |
| Oxidation Stability | ASTM D525 | minutes | 1440 |
| Copper Corrosion | ASTM D130 | 1 |
| Gum content, washed | ASTM D381 | mg/100ml | 5 |
| Research Octane Number | ASTM D2699 | 96.0 | 97.5 |
| Motor Octane Number | ASTM D2700 | Report |
| R+M/2 | D2699/2700 | Report |
| Sensitivity | D2699/2700 | 7.5 |
| Net Heat of Combustion | ASTM D240 | btu/lb | Report |
| Color | Visual | Green |
| Cetane Index | ASTM D4737 | 50 |
| Gross BTU Value | ASTM D240 | btu/lb | Report |
| Oxidation Stability | ASTM D2274 | mg/100ml | 2.5 |
| Lubricity, HFRR, @ 60°C | ASTM D6079 | mm | 0.400 |

FIG. A4.1 KA24E Test Fuel

AX.  SAFETY PRECAUTIONS

A5.1  General Information:

A5.1.1  The operating of engine tests can expose personnel and facilities to a number of safety hazards. It is recommended that only personnel who are thoroughly trained and experienced in engine testing should undertake the design, installation, and operations of engine test stands.

A5.1.2  Each laboratory conducting engine tests should have their test installation inspected and approved by their safety department. Personnel working on the engines should be provided with proper tools, be alert to common sense safety practices, and avoid contact with moving, or hot engine parts, or both. Guards should be installed around all external moving or hot parts. When engines are operating at high speeds, heavy-duty guards are required, and personnel should be cautioned against working alongside the engine and coupling shaft. Barrier protection should be provided for personnel. All fuel lines, oil lines, and electrical wiring should be properly routed, guarded, and kept in good order. Scraped knuckles, minor burns, and cuts are common, if proper safety precautions are not taken. Safety masks or glasses should always be worn by personnel working on the engines and no loose or flowing clothing, including long hair or other accessory to dress which could become entangled, should be worn near running engines.

A5.1.3  The external parts of the engines and the floor area around the engines should be kept clean and free of oil and fuel spills. In addition, all working areas should be free of tripping hazards. Personnel should be alert for leaking fuel or exhaust gas. Leaking fuel represents a fire hazard and exhaust gas fumes are noxious. Containers of oil or fuel cannot be permitted to accumulate in the testing area.

A5.1.4  The test installation should be equipped with a fuel shutoff valve, which is designed to automatically cut off the fuel supply to the engine when the engine is not running. A remote station for cutting off fuel from the test stand is recommended. Suitable interlocks should be provided so that the engine is automatically shut down when any of the following events occur: engine loses oil pressure, dynamometer loses field current, engine overspeeds, exhaust system fails, room ventilation fails, or the fire protection system is activated.

A5.1.5  Consider an excessive vibration pickup interlock if equipment operates unattended. Fixed fire protection equipment should be provided.

A5.1.6  Normal precautions should be observed whenever using flammable solvents for cleaning purposes. Make sure adequate fire fighting equipment is immediately accessible.

ASTM International takes no position respecting the validity of any patent rights asserted in connection with any item mentioned in this standard. Users of this standard are expressly advised that determination of the validity of any such patent rights, and the risk of infringement of such rights, are entirely their own responsibility.

This standard is subject to revision at any time by the responsible technical committee and must be reviewed every five years and if not revised, either reapproved or withdrawn. Your comments are invited either for revision of this standard or for additional standards and should be addressed to ASTM International Headquarters. Your comments will receive careful consideration at a meeting of the responsible technical committee, which you may attend. If you feel that your comments have not received a fair hearing you should make your views known to the ASTM Committee on Standards, at the address shown below.

This standard is copyrighted by ASTM International, 100 Barr Harbor Drive, PO Box C700, West Conshohocken, PA 19428-2959, United States. Individual reprints (single or multiple copies) of this standard may be obtained by contacting ASTM at the above address or at 610-832-9585 (phone), 610-832-9555 (fax), or service@astm.org (e-mail); or through the ASTM website (www.astm.org). Permission rights to photocopy the standard may also be secured from the ASTM website (www.astm.org/COPYRIGHT/).

# Footnotes

This test method is under the jurisdiction of ASTM Committee D02 on Petroleum Products, Liquid Fuels, and Lubricants   and is the direct responsibility of Subcommittee D02.B0 on Automotive Lubricants.

Current edition approved May 1, 2014. Published June 2014. Originally approved in 2003. Last previous edition approved in 2013 as D6891 – 13a. DOI: 10.1520/D6891-14.

The ASTM Test Monitoring Center will update changes in this test method by means of Information Letters. Information letters may be obtained from the ASTM Test Monitoring Center (TMC), 6555 Penn Ave., Pittsburgh, PA 15206-4489, Attention: Administrator. www.astmtmc.cmu.edu. This edition incorporates all Information Letters through No. 13–1.

For referenced ASTM standards, visit the ASTM website, www.astm.org, or contact ASTM Customer Service at service@astm.org. For  Annual Book of ASTM Standards volume information, refer to the standard’s Document Summary page on the ASTM website.

The last approved version of this historical standard is referenced on www.astm.org.

Available from The American Petroleum Institute (API), 1220 L. St., NW, Washington, DC 20005.

Available from Society of Automotive Engineers (SAE), 400 Commonwealth Dr., Warrendale, PA 15096-0001.

Available from American Society of Mechanical Engineers (ASME), ASME International Headquarters, Three Park Ave., New York, NY 10016-5990.

Available from Japanese Standards Organization (JSA), 4-1-24 Akasaka Minato-Ku, Tokyo, 107-8440, Japan.

Available from the Coordinating European Council for the Development of Performance Tests Transportation Fuels, Lubes, and other Fluids, Madou Plaza , 25 Floor Place, Madou B-1210, Brussels, Belgium.

Available from ASTM Test Monitoring Center (TMC), 6555 Penn Avenue, Pittsburgh, PA 15206-4489, Attention: Administrator.

The sole source of supply of the apparatus known to the committee at this time is Lambda Electronics Corporation, 515 Broad Hollow Road, Melville, NY 11747-3700.

The sole source of supply of the apparatus known to the committee at this time is Dowell Chemical Company, 1201 South Sheldon Road, Channelview, TX 77530-0429.

The sole source of supply of the apparatus known to the committee at this time is Micromotion, 7070 Winchester Circle, Boulder, CO 80301.

The sole source of supply of the apparatus known to the committee at this time is Horiba Instruments, 17671 Armstrong Avenue, Irvine, CA 92714.

The sole source of supply of the apparatus known to the committee at this time is ETAS, 2155 Jackson Avenue, Ann Arbor, MI 48103.

The sole source of supply of the apparatus known to the committee at this time is Texaco Lubricants Company, P.O. Box 4427, Houston, TX 77210-4427.

The sole source of supply of the apparatus known to the committee at this time is Barco, Hyspan Precision Products, 1685-T Brandwine Avenue, Chula Vista, CA 91911.

The sole source of supply of the apparatus known to the committee at this time is Sierra Instruments, 5 Harris Court, Monterey, CA 93940.

The sole source of supply of the apparatus known to the committee at this time is Badger Meter, Inc., Precision Products Division, 6116 East 15th Street, Tulsa, OK 74112.

Supporting data have been filed at ASTM International Headquarters and may be obtained by requesting Research Report RR:D02-1218.

The sole source of supply of the apparatus known to the committee at this time is Precision Devices, Inc., 606 County Street, P.O. Box 220, Milan, MI 48160-0220.

1994 Nissan Service Manual, SM4E-0D21U0, Model D21 Series, printed August 19, 1993, available from the local Nissan dealership’s service department.

American Chemistry Council, 1300 Wilson Boulevard, Arlington, VA 22209.

# Text boxes

Oberg Filter Housing